## WHO_MDA_ digital xray_techspecs
Annex 1
ITB 2024/EURO/MDA/0004
Digital radiographic system, pediatric
NAME, CATEGORY AND CODING
Parameters |  |  | Minimum expected Specification
1
2 | Generic name | X-ray system
3 | Specific type or variation | Stationary, digital
4 | GMDN name | Stationary basic diagnostic x-ray system, digital
5 | GMDN code | 37645
6 | GMDN category | Diagnostic and therapeutic radiation devices
7 | UMDNS name | Radiographic/Fluoroscopic Systems, General-Purpose | UMDNS - Universal Medical Device Nomenclature System For more information and downloads see the website on UMDNS.
8 | UMDNS code | 16-885
9 | Alternative name | Basic radiologic system (BRS); Digital radiography; Electronically recorded digital radiography; General radiographic x-ray equipment; Radiographic unit, chest; Radiographic unit, general-purpose; pediatric radiographic unit; Radiographic unit, skeletal
10 | Keywords | imaging, radiology
11 | GMDN definition | An assembly of devices that comprise general-purpose stationary diagnostic x-ray system used in a variety of routine planar x-ray imaging applications. It uses digital techniques for image capture, display and manipulation. The stationary design requires it to be installed and used in a fixed location within a building or in a transportation van (a mobile imaging van). This system consists of modular configurations that can be upgraded by the addition of hardware/software components.
PURPOSE OF USE
1 | Clinical or other purpose | Intended in pediatric diagnostic with radiographic examinations.
2 | Level of use | District hospital, provincial hospital, national hospital, specialized hospital.
3 | Clinical department/ward | Intensive Care Unit, Emergency, Recovery, Medical Imaging, Outpatient Services, Inpatient Care, Specialized Diagnostics, Specialized Treatment
4 | Overwiew of functional requirements | Provides X-ray film images of all parts of the body X ray generator and image intensifier can be moved to image required body part Control unit to be separate for operation from behind protective screens DICOM compatible image storage and transfer required
5 | Radiology technology | Purchase in series | optional
|  | Longitudinal scanning type acquisition with slot | optional
|  | Acquisition mode for reconstructed radiography scan | required, Automatic or Manual
|  | Automatic scan radiography reconstruction on the acquisition station | required, Automatic or Manual
|  | Preset programs with the possibility of manual modification for different anatomical areas | min. 3
|  | Scan area | min. 42x42 cm
|  | Post processing | required
|  | Zoom/Zoom level control; Rotate/Reverse control | required
|  | Annotations on image | marker, text, lines, angle, missing
TECHNICAL CHARACTERISTICS
6 | General configuration | X ray Tube above the table | required
|  | Controlled table | required
|  | Control panel in the control room | required
|  | Intercom device: Patient-doctor communication (control room) | required
7 | X-ray generator | Output power | ≥50 kW
|  | Radiography regime kV range | ≥40-150kV
|  | Radiographic range | ≥10 - 600 mAs
|  | x-ray output current | ≥600 mA
|  | AEC (Automatic exposure control) | ≥3 fields
8 | X-ray tube | Thermal storage capacity of the assembly | ≥500 kHU
|  | Tube rotation | ± 90°
|  | Focal spot | ≥ 2
|  | Small focal spot, max. size. | 0.6mm
|  | Large focal spot, size | ≥1.2 mm
|  | kV, x-ray | ≥150kV
|  | Tube overload protection | required
|  | Automatic mains voltage compensation | required
|  | Instalation/fixing | on the floor
|  | Rotating type tube | required
|  | Anode speed variable, | up to 10000rot/min
9 | X-ray Bucky Stand | Size, minimum | 43x43 cm
|  | Automatic Exposure Control (AEC) | required min. 3 fields
|  | Grid ratio | ≥10:1
|  | Box size | ≥43x43
|  | Electromagnetic control system | required
|  | Longitudinal displacement of x-ray tub suport , cm | min. 150
10 | Examination table | sincronized with tube | required
|  | Patient weight | ≥150 kg
|  | Patient weight when reclining new value - removed | removed
|  | Dimensions | ≥220x70 cm
|  | Lateral movement, both sides | ≥15 cm
|  | Longitudinal movement | ≥ + - 50 cm
|  | Vertical motorized | required
|  | Table density | < 1 mm Al or equivalent
|  | Position indicators , new value-removed | removed
|  | Speed adjustment for motorized directions in steps: new value - removed | removed
11 | The image tower
|  | Removable grid | required
|  | source image receptor distance (SID) max. | ≥150 cm
|  | source image receptor distance (SID) min. | 90 - 110 cm
|  | Control panel, of the generator in the control room. | required
|  | Control panel for tube and table position in the examination room | required
12 | Detector | type of detector | Digital Flat Detector
|  | number of DFD | 2
|  | Pixel depth | min 16 bit
| Table flat digital detector | type of conection detector | wireless or wired
|  | Size, minimum | ≥43x43 cm, (17” x 17”)
|  | Matrix resolution, pixels | ≥ 3032x3032 pixels
|  | DQE @ ≤ 1.0 lp/mm | ≥ 65 %
|  | Pixel size | ≤140 µm
| Bucky vertical flat digital detector | type of detector | wireless
|  | Size, minimum | ≥35x43 cm, (14” x 17”)
|  | Matrix resolution, pixels | ≥ 2540x3032 pixels
|  | DQE @ ≤ 1.0 lp/mm | ≥ 65 %
|  | Pixel size | ≤140 µm
13 | COLLIMATOR | Working regime min. | automatically or manually
|  | light source type | LED
|  | Collimation type | min. rectangular
|  | Built-in filters | min. 3 types
|  | Filter selection | Automatic or manual
|  | Screen | required
|  | Motorized | required
|  | Display | SID, Filter type, inclination
|  | Synchronization | required
14 | PACS server hardware, minimal configuratrion | the server must be installed, tested, synchronized with the radiological installation, accessed through the institution's network, with all necessary accessories present | required
|  | Processor | min. 6 cores, min. 2.5 GHz equivalent or better performance
|  | HDD | Minimum 2 TB of usable storage, RAID
|  | RAM | ≥ 16GB
|  | monitor | 1 pcs.
|  | monitor size | min 21 inches
|  | lum.LED | min 400 CD/m2
|  | monitor resolution | min 1920×1080
|  | installed PACS systems | Yes
|  | Accessory | Mouse, keyboard
|  | number of licences | min 5
15 | Working station/universal withradiologist | Processor | suitable for operating a x-ray system and DICOM vewer
|  | SSD/HDD | NVME /M2 /SATA3, min. 512 Gb
|  | RAM | ≥ 16GB
|  | Medical monitor | minim 2 pcs
|  | Medical monitor | min 24 inches
|  | Medical monitor | min 400 CD/m2
|  | resolution | min 1,920 x 1,080 pixels
|  | CD/DVD | Yes
|  | Compatible with PACS systems | Yes
|  | USB | ≥ 4 ports
|  | Accessory | Mouse, keyboard
|  | X-ray | Targeted images, "4in one"
|  | custom presets | required
|  | DICOM package | installed
16 | Other accessories | Patient radiation protection set (1 apron, 1 thyroid gland protection collar) for children -min 0,35 mm pb | required
|  | Operator radiation protection set (apron) | 2 sets. (min 0,35 mm Pb)
|  | Printer (Thermal direct ) spatial resolution ≥12, Format film 20x25-35x43, DICOM interface, Avalabile simultanious format ≥2), 1 pcs | required
|  | Thermal film 20x25 (500)-35x43 (500) | 1000 pcs
|  | DAP-meter integrated, with display of results in the database of the patient | required
|  | Printer, laser, for printing reports, A4 format, 1pcs | required
|  | Child restraint systems | 2 sets
17 | Software | All software and user licences must be with lifetime terms | required
|  | Compatible with hardware for radiography and fluoroscopy | included
|  | The minimum options of DICOM viewer | required
|  | - DICOM Verification | included
|  | - DICOM Modality worklist | included
|  | - DICOM MPPS | included
|  | - DICOM Storage | included
|  | - DICOM Storage commitment | included
|  | - DICOM Grayscale print | included
|  | - DICOM Query / Retrieve | included
|  | - DICOM GSPS | included
|  | import/export data | included
|  | export data on DVD/CD | included
|  | PACS | required
Licenses will be presented in electronic format or on paper with an unlimited validity period.
Image manipulation functions
Measurements, annotations and saving.
High-resolution multi-monitor diagnostic display
18 | Image System & Connectivity, other | Image storage | minim 2 Tb server and minim 512Gb on working station
|  | Processing imported images | DICOM 3.0 compatible
|  | Export images | USB
UTILITY REQUIREMENTS
19 | Electrical | The voltage of the electric current | 220, 380 V ±10%; 50 Hz
|  | Plug type in compliance with required local/national standards and regulations | required
|  | Appropriate external device to protect the equipment against over-voltage and over-current line conditions (between plug and socket) | required
|  | UPS included, installed and sinhronized with digital side of unit. Must have 20 minuts of autonomy. | required
TRAINING, INSTALLATION AND UTILISATION
20 | Transportation | Supplier must include the transportation to the final healthcare facility | required
21 | Installation | Supplier must perform the installation, safety and operation checks before handover | required
22 | Training of user/s | Training of users in hardware and software (including applications): | required
|  | • Basic operation of the system | required
|  | • Advanced operation of the system | required
|  | • Maintenance routine | required
WARRANTY AND MAINTENANCE
23 | Warranty | minim 24 months | required
24 | Preventive maintenance | Included in the warranty according with the interval and period recommended by the manufacturer. All nececary consumables and materials needed for maintenance during the given period to be included in the offer. | required
25 | Software / Hardware upgrade availability | (any avaiable updated during the warranty period) Included | required
DOCUMENTATION
26 | Documentation requirements | All supporting documentation, Technical/Maintenance manuals to be supplied in English. User manual in romanian language, according with Moldovan law nr. 102 about medical devices including: | required
|  | • Manual procedures required for local calibration and routine maintenance | required
|  | • Troubleshooting protocols | required
|  | Full set of requirement documents for registering unit on agency of National Agency for the Regulation of Nuclear and Radiological Activities | required
|  | • Cleaning and disinfection protocols | required
|  | List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Certificate of calibration and inspection to be provided. Contact details of manufacturer, supplier and local service agent to be provided | required
SAFETY AND STANDARDS
27 | Regulatory Approval / Certification | Free Sales Certificate (FSC) provided by any of the following countries: Australia, Canada, Japan, USA and/or European Community (e.g. CE certificate given by a third certified party for the specific medical devices proposed (not only declaration of conformity)). | required
28 | Standards for the manufacturer | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60336:2005 (X-ray tube assemblies for medical diagnosis - Characteristics of focal spots) IEC 60522:1999 (Determination of the permanent filtration of X-ray tube assemblies) IEC 60526:1978 (High-voltage cable plug and socket connections for medical X-ray equipment) IEC 60601-1-3:2013 (Part 1-3: General requirements for basic safety and essential performance - Collateral Standard:Radiation protection in diagnostic X-ray equipment) IEC 60601-2-28:2010 (Part 2-28: Particular requirements for the basic safety and essential performance of X-ray tube assemblies for medical diagnosis) IEC 60601-2-43:2010 (Part 2-43: Particular requirements for the basic safety and essential performance of X-ray equipment for interventional procedures) IEC 60601-2-54:2009 (Part 2-54: Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and radioscopy) IEC 60613:2010 (Electrical and loading characteristics of X-ray tube assemblies for medical diagnosis) EN 62304 Software for medical devices. Software life cycle processes (IEC 62304:2006) EN 62304:2006/AC:2008 | required