## Shaver system for orthopaedics
Technical Responsiveness Table
Please indicate Name of the Company/Bidder
Code No. |  | - | Shaver system for orthopaedics |  | Quantity: |  |  | 1 set |  | Please indicate Brande name/Model
Purpose : | [9 empty] | Please provide below exact data as per the offered model
Universal device console for connecting power-assisted shaver, drill and saw tools.
1. COMPOSITION
1-1 | Composition
| (1) | Shaver console |  | : | 1 unit | x | 1 | = | 1 unit
| (2) | Handpiece |  | : | 1 unit | x | 1 | = | 1 unit
| (3) | Foot switch |  | : | 1 unit | x | 1 | = | 1 unit
1-2 | Accessories
Any additional component or accessory which is required for adequate installation and operation should be included and described on your offered specification.
2. SPECIFICATIONS
2-1 | Motor system for orthopaedics
| (1) | Shaver console |  | : | *Equipped
| (2) | Shaver allows clock and counterclockwise rotations |  | : | Equipped
| (3) | Shaver clockwise rotation |  | : | min 8000 rpm
| (4) | Shaver counterclockwise rotation |  | : | min 8000 rpm
| (5) | Shaver' modes of oscillations |  | : | 7 or more
| (6) | Shaver blade diameter |  | : | **from 2 to 8 mm at least 5 different sizes (minor deviations expected not all have 5 sizes in multiple use)
| (7) | Clockwise, counterclockwise and oscillation handpiece rotation modes |  | : | Equipped
| (8) | Ergonomic and autoclavable handpiece |  | : | Equipped
| (9) | Automatic detection of a connected handpiece |  | : | Equipped
| (10) | Handpiece connection to control unit |  | : | wired or wireless
| (11) | The shaver system can be controlled by a footswitch |  | : | Equipped
| (12) | Memory storage |  | : | Equipped
| (13) | Handpiece equipped with 2 buttons, each having double functionality |  | : | Equipped
| (14) | Power requirements |  | : | AC 220V±10%, single phase, 50Hz
3. | QUALITY ASSURANCE REQUIREMENTS
| (1) | EC Certification
| (2) | Compliant to EN 60601-1 or 60601-2 and other applicable EU standards
| (3) | Manufacturer ISO 13485
| (4) | Declaration of conformity with guideline 93/42/EEC
| (5) | EN 60529, IPX (7,8) testing for handpiece and footswitch |  | : | (console is not requiring IPX standard since it is not expected to be held in liquids)
| (6) | Warranty at least 12 months
4. | SPARE PARTS
Not required
5. | CONSUMABLES
| (1) | Reusable shaver blades |  | : | ≥ 5
6. | NOTE
*Console compatible with the handpiece and foot switch.
**Selection of the blades and endings (resectors, burrs or cutters) will be made by the beneficiary after awarding the bidder
Equipment assembly, installation and commissioning on site at beneficiary location
User’s manuals, preferable in Romanian or Russian language (optional English)
Maintenance manual, preferable in Romanian or Russian language (optional English)