## Arthroscope
Technical Responsiveness Table
Please indicate Name of the Company/Bidder
Code No. |  | - | Arthroscope |  | Quantity: |  |  |  | 1 set | Please indicate Brande name/Model
Purpose : | [9 empty] | Please provide below exact data as per the offered model
Equipment for minimally invasive for joints surgeries.
1. COMPOSITION
1-1 | Composition
| (1) | Telescope |  | : | 1 unit | x | 1 | = | 1 unit
| (2) | Trocar with sheath |  | : | 1 pc | x | 1 | = | 1 pc
| (3) | Light source |  | : | 1 unit | x | 1 | = | 1 units
| (4) | Fiber optic light cable |  | : | 1 unit | x | 2 | = | 2 unit
| (5) | Video camera head |  | : | 1 unit | x | 2 | = | 2 units
| (6) | Video camera control unit |  | : | 1 unit | x | 1 | = | 1 unit
| (7) | HD Monitor |  | : | 1 unit | x | 1 | = | 1 unit
| (8) | Trolley |  | : | 1 unit | x | 1 | = | 1 unit
1-2 | Accessories
Any additional component or accessory which is required for adequate installation and operation should be included and described on your offered specification
2. SPECIFICATIONS
2-1 | Telescope
| (1) | Type |  | : | Autoclavable
| (2) | View direction |  | : | 30̊
| (3) | Diameter |  | : | 4 mm
| (4) | Working length |  | : | 140-180 mm
| (5) | Lighting |  | : | Fiber optic transmission units
2-3 | Trocar with sheath
| (1) | Trocar |  | : | applicable for 2-1
| (2) | Diameter |  | : | 6.0 mm
| (3) | Working length |  | : | 130 mm
2-4 | Light source
| (1) | Type of light source |  | : | LED
| (2) | Light transmission method |  | : | Fiber optic cable
2-5 | Fiber optic light cable
| (1) | Diameter |  | : | 3.5 mm or more
| (2) | Length |  | : | 200 cm or more
2-6 | Video camera head
| (1) | Resolution |  | : | 1080p or higher
| (2) | Type |  | : | Autoclavable
| (3) | Controls |  | : | 2 programable control buttons
2-7 | Video camera control unit
| (1) | Resolution |  | : | 1080p or higher
| (2) | Light source |  | : | LED
| (3) | Adapters |  | : | For 4 different types of lighting cables
| (4) | USB |  | : | Equipped
| (5) | Fiber optic light transmission |  | : | Equipped
2-8 | HD Monitor
| (1) | Size |  | : | 26" or more
| (2) | Resolution |  | : | 1080p or more
2-9 | Trolley on castors
| (1) | Application |  | : | For installation of light source, video camera
control unit, image management unit, LCD
monitor, electrosurgical generator
and irrigation pump or more
| (2) | Integrated electrical sockets |  | : | Equipped
| (3) | Mobile support for the power suction |  | : | Equipped
| (4) | Earth leakage monitor or similar current protection tool |  | : | Equipped
| (5) | Antistatic castors with brakes |  | : | Equipped
3. | QUALITY ASSURANCE REQUIREMENTS
| (1) | EC Certification
| (2) | Manufacturer ISO 13485
| (3) | Declaration of conformity with guideline 93/42/EEC
| (4) | Warranty at least 12 months
4. | SPARE PARTS
4-1 | Not required
5. | CONSUMABLES
5-1 | Not required
6. | NOTE
Equipment assembly, installation and commissioning on site at beneficiary location
User’s manuals, preferable in Romanian or Russian language (optional English)
Maintenance manual, preferable in Romanian or Russian language (optional English)
Compatible with suction-irrigation unit (same manufactrurer)