## Ortopantomograph
Tehnical Responsiveness Table
Lot 1: Ortopantomograph | [5 empty] | Please indicate Name of the Company/Bidder
Purpose: A system of extraoral radiography for digital panoramic radiographs, indicated for the production of orthopanoramic images with the possibility of exposure on sections of the maxillofacial area, diagnostic examination of the tooth (teeth), arches and other structures of the oral cavity. |  |  |  | Q-ty: | 1 | Please indicate Brande name/Model
Please provide below exact data as per the offered model
1. | Specifications for Anode
- | Tube voltage | : | 60-95 kV
- | Tube current | : | 1-16 mA
- | Target angle | : | 5 °
- | Focal spot | : | 0.5 (IEC 60336/2005)
- | Total filtration | : | minimum 3.2 mm Al. at 95 kV
- | Electrical connections
- | Nominal voltage of main network | : | 220-240V
- | Tolerance | : | ± 10%
- | Input power frequency | : | 50/60 Hz
- | Nominal current rating | : | 10 A la 220–240 V
- | Main fuses (F1 and F2) | : | 220–240V
- | Power consumption | : | <2.3 kVA at 220–240V
- | Exposure controller cable, remote exposure switch, maximum 15 m | : | equipped
1.1. | Specifications for Generator
- | High frequency DC Generator | : | yes
- | Anode voltage | : | 60–95 kV (+/- 5 kV)
- | Anode current | : | 2–16 mA (+/- 20%)
- | Exposure accuracy | : | +/– 5%
- | DAP display accuracy | : | +/– 25%
- | Spine compensation compensated kV / mA | : | yes
- | Accurate amount of radiation emitted | : | > 95%
1.2. | Specifications for 2D image detector
- | Sensor technology / type | : | CMOS
- | Image reception area | : | yes
- | PAN | : | 152 x 6 mm
a) | Exposure programs | :
- | Standard panoramic program | : | <9.0 s
- | Standard segmented panoramic program | : | 1,4–9,0 s
- | Pediatric panoramic program | : | <9.0 s
- | Pediatric segmented panoramic program | : | 1,4–9,0 s
- | Bitewing | : | <5,20 s (2,6 s + 2,6 s)
- | Temporomandibular joint (TMJ), lateral projection | : | <4,0 s (2,0 s + 2,0 s)
- | Segmented panoramic program | : | (1-5) quadrants
b) | Optimization programs
- | AIE (Automatic Image Processing) | : | yes
- | Protection against electric shock | : | Class I
- | Degree of protection | : | Type B applied without conductive connection to the patient
- | Protection against external influence | : | Min. IP20
1.3. | Specifications for PC
- | Monitor | : | 22 Inch
- | Processor | : | i5-i7
- | Virtual memory | : | 4-8 Gb
- | Video card | : | 2-4 Gb, 256 bit, (min)
- | Power supply | : | min 500w
- | Hard Disk 1 | : | 500 Gb
- | Hard Disk SSD | : | 256 Gb
- | Keyboard | : | equipped
- | Mouse | : | equipped
1.4. | Specifications for Printer | :
- | Color printer | : | yes
- | Power supply | : | 220V-240V, 50-60Hz
- | Nominal current rating | : | 0,3 A
- | Connection | : | USB and WIFI
- | Maximum resolution, dpi | : | 5760x1440
- | Format paper, max | : | A4
- | Paper density | : | 64-300 g / m2
- | Number of paint containers | : | 6