## Autoclave
Tehnical Responsiveness Table
Lot 2: Autoclave | [5 empty] | Please indicate Name of the Company/Bidder
Automatic, microprocessor-controlled steam sterilizer designed for total elimination and/or inactivation of microorganisms off medical devices and related products, using pressurized steam |  |  |  | Q-ty: | 1 | Please indicate Brande name/Model
Please provide below exact data as per the offered model
1. | Composition
- | Frontal loading | : | yes
- | Automatic with Pre-vaccum | : | yes
- | Heating (steam pulsating) | : | yes
- | Sterilization (holding) | : | yes
- | Post vacuum (drying) | : | yes
2. | Specifications
- | 4 vacuum pulses are compulsory for efficient air removal, at vacuum stage | : | yes
- | 3-phases power supply | : | 380-400V, 50/60Hz
- | All the processes to be verified through a microproccesor | : | yes
- | Featured with digital full color touch screen display, for selecting and viewing the following parameters:
a) | selected cycle | : | yes
b) | temperature | : | yes
c) | working pressure | : | yes
- | Sterilization chamber volume | : | 90 liters (±5%)
- | Chamber dimensions | : | ≥ 400 x 750 mm (Inner Ø x Depth)
- | Round shape sterilization chamber | : | yes
- | Made of stainless steel 316L | : | yes
- | The inner surface of the chamber must be sandblasted (AC-450/470 standard)/ electropolished (premium units and AC-575/5100) | : | yes
- | Equipped with water purification system for steam generator and automatic water filling | : | yes
- | Polymer foam insulation materials to ensure low thermal losses. (Insulation should not be made of Rockwool or glass rock material) | : | yes
- | Equipped with built-in steam generator | : | yes
- | Equipped with built-in ink printer for recording sterilization cycles | : | yes
- | Equipped with RS232 port for PC connection | : | yes
- | Equipped with the following security features:
a) | the door to remain locked throughout the sterilization cycle | : | yes
b) | certified safety valve to ensure optimal pressure during the sterilization process | : | yes
c) | electronic control of the pressure in the sterilization chamber. When the chamber pressure is exceeded, the cycle be stopped automatically and to activate the alarm. | : | yes
d) | built-in protections against over-voltage conditions | : | yes
e) | automatic ring type locking mechanism | : | yes
- | Display for operational cycle data and alert messages, with audio and visual indications | : | yes
- | Must have access to parameters for supervisors, engineer and operator’s level | : | yes
- | Must be equipped with energy saving mode when the unit is in stand by mode | : | yes
- | Contains minimum 5 preprogrammed sterilization cycles, with temperature range of | : | ≥ 121-134 °C
- | Maximal working pressure | : | ≥ 2.7 bar
- | Minimum working pressure | : | approx. - 0.9 bar
- | Possibility to programme personalized cycles | : | yes
- | Contain at least the following preset test programs:
a) | Bowie-Dick Test for checking the steam penetration in materials during the sterilization process | : | yes
b) | Vacuum Test for testing the leak effectiveness of the sterilization chamber | : | yes
- | Equipped with movable stainless steel internal trays/plate (minimum 2) for load | : | yes
- | Peak sound level | : | <70 dB
3 | Optional:
- | Equipped with SD card reader to log full cycle data | : | yes
- | Can be mounted on a stainless steel stand with storage space | : | yes