## Annex 4
ANNEX 4: TECHNICAL COMPLIANCE TABLE
Hyperbaric oxygen chamber (Sealed chamber in which a high-pressure environment is used primarily to treat various illnesses.) |  | : | 1 unit |  |  |  |  | Please indicate Name of the Company/Bidder
Please indicate Brande name/Model
Please provide below exact data as per the offered model
1. | Composition
(1) | Hyperbaric chamber | : | 1 unit | x | 1 | = | 1 unit
(2) | O2 concentrator | : | 1 unit | x | 2 | = | 2 pcs
(1) | AC system (Recirculating chiller) | : | 1 unit | x | 1 | = | 1 unit
1-2 | Accessories | [6 empty] | Accessories
| Any other accessory which is required for installation and operation should be included and described on your offered specification | [6 empty] | Please mention, if any
2. | Specifications
Chamber Body | [7 empty] | Chamber Body
─ | Non- inflatable, hard type hyperbaric chamber | : | yes
─ | Diameter | : | 90 ± 1 cm
─ | Cabin length | : | 230 ± 10 cm
─ | Body material | : | aluminium or other
─ | Window | : | polycarbonate or similar
─ | Internal volume | : | not less than 1400 L
─ | Supported patient weight | : | not less than 120 kg
─ | Number of places | : | 1
─ | Pressure | : | min 50 kPA or 0.5 bar
─ | Protection | : | Class II, IP44 or other
─ | Resistant to corrosion | : | yes
─ | Safety | : | power failure, overpressure
─ | Voltage | : | 220-230 V/50 Hz
─ | Show errors/alarms/alerts | : | mandatory: power failure, low and high pressure, high temperature, low O2 concentration and other
─ | Power | : | maximum 600-800 VA
Chamber operation | [7 empty] | Chamber operation
─ | Mirrored patient controls | : | minimum Start/Stop and other
─ | Interphone | : | yes
─ | Compression and decompression speed | : | yes
─ | Automatically controlled pressure system | : | yes
─ | Chamber temperature | : | yes
─ | Session time | : | yes
─ | Wifi connectivity | : | yes, (optional)
O2 Concentrator | [7 empty] | O2 Concentrator
─ | O2 flow rate | : | 2–10 LPM
─ | Oxygen purity | : | 2-9 LPM: not less than 89% and for 10 LPM: not less than 87%
─ | Noise level | : | maximum 60 dB(A)
─ | Filter set | : | at least 1
─ | Operating environment | : | 5°C to 40°C and up to 95% humidity
AC System (Recirculating chiller) | [7 empty] | AC System (Recirculating chiller)
─ | Temperature fluctuations | : | maximum ± 2°C
─ | Temperature range | : | 10 - 30°C or wider
─ | Ambient temperature and humidity | : | 10 - 35°C or wider; Humidity: 30 to 70% or wider
─ | Noise level in standby mode | : | maximum 60 dB(A)
3. | General Requirements | [6 empty] | General Requirements
─ | Declaration of conformity | : | yes
─ | Manufacturer ISO 13485 | : | yes
─ | Manufacturer's Authorization Form | : | yes
─ | Authorized Service Center in the R. of Moldova or neighbouring countries | : | yes
─ | Equipment assembly, installation and commissioning on site at beneficiary location | : | yes
4. | Warranty | [6 empty] | Warranty
─ | Hyperbaric chamber (minimum 12 months from the commissioning) | : | yes
─ | O2 concentrators (minimum 12 months each unit from the commissioning) | : | yes
─ | AC system (minimum 12 months from the commissioning) | : | yes
5. | Training | [6 empty] | Training
─ | Training for operator and biomedical engineer (maintenance protocol) | : | yes
─ | Training in Romanian or Russian language | : | yes
6. | Documentation | [6 empty] | Documentation
─ | User’s manuals for all devices, preferable in Romanian or Russian language (optional English) | : | yes
─ | Maintenance manual for all devices, preferable in Romanian or Russian language (optional English) | : | yes