## Sheet1
| WHO ITB 2022/EURO/MDA/0010 | Procurement of devices, reagents and consumables for the National Agemcy for Public Health
Annex 1: List of requirements and specifications
#No | Item description | Technical Specification | UM | Quantity | Comments
LOT: 1 Equipment for diagnostic procedures
1 | Centrifuge | Reach up to 20,000 x g in just 16 seconds. Speed range of 100-15000 rpm, adjustable in increments of 100. Fixed angle rotor that can hold 24 x 1.5/2 ml tubes. Max speed (Angle rotor): 15000 rpm. Max RCF: 20375 x g. Power supply: 120 V. Dimensions (d x w x h) : 280 x 360 x 250 mm Weight without rotor: 17 kg. Max Acceleration time: 16 s. Max Deceleration time: 18 s Calibration certificate; CE, IVD Warranty period - 36 months; The seller bears all costs of installation and testing of the machine; Installation, training and maintenance shall be performed by a qualified engineer. Upon delivery, the machine must be accompanied by the original user manual and in the state or Russian language. | pcs | 5 | Please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates for all items
2 | Microcentrifuge | Speed control range 1000–3500 rpm (increment 100 rpm). Spin timer 1 s – 99 min. Rotor for 12 x 1.5/2 ml microtest tubes Calibration certificate, CE, IVD, warranty period - 36 months; The seller bears all costs of installation and testing of the machine; Installation, training and maintenance shall be performed by a qualified engineer. Upon delivery, the machine must be accompanied by the original user manual and in the state or Russian language. | pcs | 4
3 | Combi spin with minicentrifuge and vortex | Centrifugation speed 2.400 rpm at 500 x G Vortex speed 2.400 rpm Protective lid and autostop, when lid is open Continuous and impulse operation Weight: 1,7 kg, Dimensions are 190 x 235 x 125 mm Two rotor-discs, one for 12 x 0,5 ml and 12 x 0,2 ml and one for 12 x 1,5 ml tubes Calibration certificate, warranty period - 36 months; CE, IVD The seller bears all costs of installation and testing of the machine; Installation, training and maintenance shall be performed by a qualified engineer. Upon delivery, the machine must be accompanied by the original user manual and in the state or Russian language. | pcs | 9
4 | FTA-2i Advanced Aspirator with Trap Flask | Aspiration speed - up to 10 l/min (air); Trap flask -2 l, polypropylene (autoclavable); Liquid level sensor type ; Vacuum -200 to -800 mbar (adjustable); Overflow protection -motor stops, light and sound signal; Dimensions with trap flask (W×D×H)- 185 × 290 × 390 mm; Weight with trap flask-1.85 kg; Input current/power consumption12 V, 1 A / 10.8 W; External power supply Input AC 100-240V 50/60 Hz; Output DC 12 V Warranty period - 36 months; CE, IVD The seller bears all costs of installation and testing of the machine; Installation, training and maintenance shall be performed by a qualified engineer. Upon delivery, the machine must be accompanied by the original user manual and in the state or Russian language. | pcs | 5
5 | Freezing rack for PCR | For 96 x 0.1 or 96 x 0.2 ml tubes with snap top or screw cap. Maintains a sample temperature of less than 4°C for approx. 4 hours. Dramatic change to contrasting color at 7°C. Dimensions (rack + lid): 70 x 105 x 146 mm. Packaging unit: 2 racks with lid Color change at temperature deviations (from cold to warm) | pcs | 5
6 | Mono channel automatic pipette autoclavable 0.5-10 ul | Adjustable volume Fully autoclavable and UV-light resistant Universal cone for commercially available pipet tips Calibration certificate, 2 year warranty, CE | pcs | 11
7 | Mono channel automatic pipette autoclavable 2-20 ul | Adjustable volume Fully autoclavable and UV-light resistant Universal cone for commercially available pipet tips Calibration certificate, 2 year warranty, CE | pcs | 11
8 | Mono channel automatic pipette autoclavable 10-100 ul | Adjustable volume Fully autoclavable and UV-light resistant Universal cone for commercially available pipet tips Calibration certificate, 2 year warranty, CE | pcs | 11
9 | Mono channel automatic pipette autoclavable 50-200 ul | Adjustable volume Fully autoclavable and UV-light resistant Universal cone for commercially available pipet tips Calibration certificate, 2 year warranty, CE | pcs | 9
10 | Mono channel automatic pipette autoclavable 100-1000 ul | Adjustable volume Fully autoclavable and UV-light resistant Universal cone for commercially available pipet tips Calibration certificate, 2 year warranty, CE | pcs | 10
11 | Multichannel automatic pipette 8 channel, volume 10-100 ul | Adjustable volume Fully autoclavable and UV-light resistant Universal cone for commercially available pipet tips Calibration certificate, 2 year warranty, CE | pcs | 1
12 | ELISA-Reader | Voltage-100-240 V AC; Dimension-425 x 310 x 215 mm (LxWxH); Reading Speed - 5 Seconds For 96- Well Plate; Frequency - 50/60 Hz; Absorbance Range 0-4.000 Abs; Resolution 0.001 Abs (Displayed); 0.0001 Abs (Calculated; Light Source 6 V/ 19 W Halogen Lamp; Wavelength Accuracy +/- 2 nm; Band Width 8+/-2 nm; Fixed time, End Point methods Large LCD display with touch screen, 4 fixed filters and 3 more filters optional 8-channel optical fiber system, auto self-check when power on Powerful QC function Quality Control Data Check Function 7 inch Color LCD Touch Screen with Built-in Printer Plate Shaking Function, Speed and Frequency Adjustable Improved user experience with powerful software Effective data and result interpretation Multi-assay enables upto 12 different assays in one plate. USB port for easy data transfer Warranty period - 36 months; The seller bears all costs of installation and testing of the machine; Installation, training and maintenance shall be performed by a qualified engineer and certified by the manufacturer. Upon delivery, the machine must be accompanied by the original user manual and in the state or Russian language. | pcs | 1
13 | Vortex | Speed range -500-2500 rpm. Operating mode - touch or continuous. Speed control - 0-6 Operating mode - uninterrupted Warranty period - 36 months The seller bears all costs of installation and testing of the machine; Installation, training and maintenance shall be performed by a qualified engineer and certified by the manufacturer. Upon delivery, the machine must be accompanied by the original passport with the manufacturer's stamp and the user manual in the state or Russian language. | pcs | 13
14 | Thermoshaker | Temperature setting range +25°C ... +100°C; Temperature control range 5°C above ambient ... +100°C; Temperature setting resolution 0.1°C; Temperature stability ±0.1°C; Temperature accuracy at +37°C-±0.5°C; Average heating speed from +25°C to +100°C-4°C/min; Temperature uniformity over the block at +37°C-±0.1°C; Temperature uniformity over the block at +100°C-±0.2°C; Temperature calibration coefficient range-0.936...1.063 (± 0.063); Speed control range 250–1400 rpm; Digital time setting-1 min–96 hrs (1 min increment); Timer sound signal; Orbit 2 mm; Microprocessor controlled temperature, mixing speed and operation time; Maximum continuous operation time -168 h; Overall dimensions (W×D×H) 220x240x90 mm; External power supply Input AC 100–240 V; 50/60 Hz; Output DC. Warranty: 3 years The seller bears all costs of installation and testing of the machine; Installation, training and maintenance shall be performed by a qualified engineer and certified by the manufacturer. Upon delivery, the machine must be accompanied by the original user manual and in the state or Russian language. | pcs | 5
15 | ELISA Washing tray | Logic and intuitive user interface Liquid level sensors (wash and waste bottle) Plate sensor Includes: 1x2L wash bottle, 1x2L waste bottle, 1x8L wash head, aerosol cover Dispense Volume-500 nL to 3000 μL/well; Washing Speed- 96 wells: 300 μL/well, 96-tube manifold: 13 s; 384 wells: 100 μL/well, 192-tube manifold: 17 s; 1536 wells: 10 μL/wells, two 32-tube manifolds: 36 s; Precision- Wash: <3% CV (model dependent); Dispense (Syringe Pump): ≤5% CV at 5 μL, ≤2.5% CV at 20 μL, ≤1% CV at 100 μL; Soak Time- Programmable in minutes and seconds, up to 60 min; Residual Volume ≤2 μL/well. Warranty: 3 years The seller bears all costs of installation and testing of the machine; Installation, training and maintenance shall be performed by a qualified engineer and certified by the manufacturer. Upon delivery, the machine must be accompanied by the original passport with the manufacturer's stamp and the user manual in the state or Russian language. | pcs | 1
16 | Inoculator (similar to RETRO C80) | Inoculator for ETEST and for diffusimetric sensitivity testing, used for large (150mm) and small (90mm) plates. Speed: 0-500 rpm, Noise level: Max 60 dB (A), Dimensions: 180x180x90 mm (W x D x H), Weight 3.5 kg. Warranty: 3 years The seller bears all costs of installation and testing of the machine; Installation, training and maintenance shall be performed by a qualified engineer and certified by the manufacturer. Upon delivery, the machine must be accompanied by the original passport with the manufacturer's stamp and the user manual in the state or Russian language. | pcs | 11
17 | Device to measure the density of the inoculum (with all components), similar to Vitek Densichek | Offers an intuitive user experience and convenient touchscreen to display McFarland values. Stable McFarland standards that don't expire Broad reading range from 0.00 to 4.00 McFarland with an accuracy and a margin of +/- (6.5% + 0.06) Automatic transfer of McFarland reading to VITEK 2 software (version 9 or higher). The device contains a basic unit with a detachable optical interface, the pod. With the help of the Pod, the user can visually examine the samples with optical readings. Warranty: 3 years The seller bears all costs of installation and testing of the machine; Installation, training and maintenance shall be performed by a qualified engineer and certified by the manufacturer. Upon delivery, the machine must be accompanied by the original user manual and in the state or Russian language. | pcs | 11
18 | Digital thermometer | Temperature range -70°C to +250°C in 4-5 sec, accuracy ±0,5°C at -20…150°C, rest ±1°C, reduced tip provides faster & more accurate readings, waterproof, fast and accurate, °C/°F switchable, stainless steel probe 116 mm, for universal use, calibration certificate Warranty: 1 year | pcs | 10
19 | Electrophoresis system (similar to Invitrogen E-Gel Power Snap Electrophoresis system) | Electrophoresis system with directly connected camera DNA separation in aprox 10 minutes with dry precast e-gel agarose gels View and capture DNA samples in real time Benchtop use Blue-light transilluminator -Electrophoresis device requirements: Dimensions:245x128x72 mm (+-3 mm) Touch Screen LCD display : minimal 75x40 mm Viewing surface dimension :92X110 mm (+-3 mm) Dimensions of amber filter: 86x106 mm(+-3 mm) Blue Led :CWL 465 nm, FWHM 20 nm Led life: at least 48000 hours -Camera requirements Dimensions:260x128x150 mm (+-3 mm) Memory: at least 30 GB Touch Screen LCD display : minimal 114x85 mm Type of camera: color CMOS Gel image resolution:1600x1944,8 bits Image output: color JPG Kit of delivery must include: 1 camera device 1 electrophoresis device 48 bags of e-gels (agaraose16x1%, 16x2% , 16x4%) 1 vial e-gel 50 bp DNA Ladder 1 power adapter 1 pair of viewing glasses Warranty: 3 years The seller bears all costs of installation and testing of the machine; Installation, training and maintenance shall be performed by a qualified engineer and certified by the manufacturer. Upon delivery, the machine must be accompanied by the original user manual and in the state or Russian language. | pcs | 1
TOTAL LOT 1: Equipment for diagnostic procedures
LOT 2: Reagents for diagnostic procedures
ELISA
1 | Coxiella burnetii Phase 1 IgA qualitativ ELISA Kit | Reaction type - immunoenzymatic, for human diagnosis. Duration of the incubation period in the test reaction - up to 120 minutes. Incubation will not include the stirring process. Stable liquid ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 98%, diagnostic specificity not less than 96%. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
2 | Coxiella burnetii Phase 2 IgM qualitativ ELISA Kit | Reaction type - immunoenzymatic, for human diagnosis. Duration of the incubation period in the test reaction - up to 120 minutes. Incubation will not include the stirring process. Stable liquid ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 98%, diagnostic specificity not less than 96%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
3 | Coxiella burnetii Phase 1 IgG qualitativ ELISA Kit | Reaction type - immunoenzymatic, for human diagnosis. Duration of the incubation period in the test reaction - up to 120 minutes. Incubation will not include the stirring process. Stable liquid ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 98%, diagnostic specificity not less than 96%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
4 | Coxiella burnetii Phase 2 IgG quantitativ ELISA Kit | Reaction type - immunoenzymatic, for human diagnosis. Duration of the incubation period in the test reaction - up to 120 minutes. Incubation will not include the stirring process. Stable liquid ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 98%, diagnostic specificity not less than 96%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
5 | Enterovirus Ag ELISA Kit | Reaction type - immunoenzymatic, for human diagnosis. Duration of the incubation period in the test reaction - up to 120 minutes. Incubation will not include the stirring process. Stable liquid ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 98%, diagnostic specificity not less than 96%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
6 | Francisella tularensis IgM ELISA Kit | Reaction type - immunoenzymatic, for human diagnosis. Duration of the incubation period in the test reaction - up to 120 minutes. Incubation will not include the stirring process. Stable liquid ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 98%, diagnostic specificity not less than 96%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
7 | Francisella tularensis IgG ELISA Kit | Reaction type - immunoenzymatic, for human diagnosis. Duration of the incubation period in the test reaction - up to 120 minutes. Incubation will not include the stirring process. Stable liquid ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 98%, diagnostic specificity not less than 96%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
8 | Hepatites A IgM ELISA | Reaction type - enzyme-linked immunosorbent assay, for human diagnosis. Detection of IgM antibodies to hepatitis Adin virus serum or human plasma. Serum and reagents Ready to work Liquid Set of all reagents required for reaction, including controls 96-well plate (12 strips of 8 removable wells) Reagents for 96 investigations including controls Test sensitivity not less than 100 %, diagnostic specificity not less than 98%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | set | 2
9 | Crimean Congo Hemorrhagic Fever IgM ELISA | Reaction type - enzyme-linked immunosorbent assay for human diagnosis. Incubation period - up to 120 minutes. Incubation will not include the stirring process. Liquid set ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 96%, diagnostic specificity not less than 99%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | set | 1
10 | Bacillus anthracis IgG ELISA Kit | Reaction type - enzyme-linked immunosorbent assay for human diagnosis. Incubation period - up to 120 minutes. Incubation will not include the stirring process. Liquid set ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 96%, diagnostic specificity not less than 99%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | set | 1
11 | Leptospirosis IgM/IgG ELISA Kit | Reaction type - enzyme-linked immunosorbent assay for human diagnosis. Incubation period - up to 120 minutes. Incubation will not include the stirring process. Liquid set ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 96%, diagnostic specificity not less than 99%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | set | 2
12 | Yersinia enterocolitica/pseudotuberculosis IgM ELISA Kit | Reaction type - enzyme-linked immunosorbent assay for human diagnosis. Incubation period - up to 120 minutes. Incubation will not include the stirring process. Liquid set ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 96%, diagnostic specificity not less than 99% | set | 2
13 | Yersinia enterocolitica/pseudotuberculosis IgG ELISA Kit | Reaction type - enzyme-linked immunosorbent assay for human diagnosis. Incubation period - up to 120 minutes. Incubation will not include the stirring process. Liquid set ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 96%, diagnostic specificity not less than 99%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | set | 2
14 | Brucella spp. IgM ELISA kit | Reaction type - enzyme-linked immunosorbent assay for human diagnosis. Incubation period - up to 120 minutes. Incubation will not include the stirring process. Liquid set ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 96%, diagnostic specificity not less than 99%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | set | 2
15 | Brucella spp. IgG ELISA kit | Reaction type - enzyme-linked immunosorbent assay for human diagnosis. Incubation period - up to 120 minutes. Incubation will not include the stirring process. Liquid set ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 96%, diagnostic specificity not less than 99%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | set | 2
16 | Coxiella burneti IgM ELISA Kit | Reaction type - enzyme-linked immunosorbent assay for human diagnosis. Incubation period - up to 120 minutes. Incubation will not include the stirring process. Liquid set ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 96%, diagnostic specificity not less than 99%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | set | 2
17 | Rickettsia prowazekii IgG ELISA Kit | Reaction type - enzyme-linked immunosorbent assay for human diagnosis. Incubation period - up to 120 minutes. Incubation will not include the stirring process. Liquid set ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 96%, diagnostic specificity not less than 99%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | set | 1
18 | Rickettsia prowazekii IgM ELISA Kit | Reaction type - enzyme-linked immunosorbent assay for human diagnosis. Incubation period - up to 120 minutes. Incubation will not include the stirring process. Liquid set ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 96%, diagnostic specificity not less than 99%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | set | 1
19 | Crimean Congo Hemorrhagic Fever IgM ELISA Kit | Reaction type - enzyme-linked immunosorbent assay for human diagnosis. Incubation period - up to 120 minutes. Incubation will not include the stirring process. Liquid set ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 96%, diagnostic specificity not less than 99%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | set | 1
20 | Crimean Congo Hemorrhagic Fever IgG ELISA Kit | Reaction type - enzyme-linked immunosorbent assay for human diagnosis. Incubation period - up to 120 minutes. Incubation will not include the stirring process. Liquid set ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 96%, diagnostic specificity not less than 99%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | set | 1
21 | Tick Borne Encephalitis Virus IgM/IgG ELISA Kit | Reaction type - enzyme-linked immunosorbent assay for human diagnosis. Incubation period - up to 120 minutes. Incubation will not include the stirring process. Liquid set ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 96%, diagnostic specificity not less than 99%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | set | 2
22 | Hantavirus Hemorrhagic Fever with Renal Syndrome IgM/IgG ELISA Kit | Reaction type - enzyme-linked immunosorbent assay for human diagnosis. Incubation period - up to 120 minutes. Incubation will not include the stirring process. Liquid set ready to work. Presence in the kit of all reagents required for the reaction, including controls. 96-well plate (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 96%, diagnostic specificity not less than 99%.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | set | 2
Elisa Total
PCR
1 | Multiplex RT PCR Respiratory kit (inlcuding COVID + Influenza) | The SARS CoV-2 virus RNA amplification kit should contain at least 2 targets of interest: - at least one specific target for SARS-CoV-2, at least one target for Influenza A / B and one target for internal control. Unit of measurement: Part = test. The kit should contain all the necessary components for the PCR reaction with prepared positive and negative control. Intra-test accuracy (repeatability): CV <3%; The intertext accuracy (reproducibility): CV <5%. The PCR reaction should be performed in a single step and the amplification time should not exceed 2 hours. Detection limit - up to 500 copies / ml or equivalent in copies per reaction. Sensitivity> 98%; Specificity> 99% Stability of the opening kit for at least 30 days. Amplification reagents to be compatible with Aplied Biosystem 7500, Quant Studio 5 equipment. Contains instructions for use. EC certificate for use of IVD (for human diagnosis) in accordance with the requirements of the European Directive. | test | 2500
2 | Multiplex RT PCR Meningitis | Reaction type - Chain polymerization reaction with real time detection. For human diagnosis. Reagents for not less than 25 tests, including controls, compatible with QuantStudio 5 amplifier IVD. Detection in CSF of Escherichia coli K1, Group B Strep, Haemophilus influenzae, Listeria monocytogenes, Neisseria meningitidis, S. pneumoniae. The set will include reagents for all stages: extraction, amplification with real-time detection.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
3 | RT PCR Measles Diagnostic Kit | Reaction type - Chain polymerization reaction with real time detection. For human diagnosis. Reagents for not less than 50 tests, including controls, compatible with QuantStudio 5 amplifier. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 2
4 | RT PCR HIV Kit | Reaction type - chain polymerization reaction (PCR) with real-time detection. (dual targeting prevents detection failure caused by possible mutations inside the HIV-1 genome. Detects a broad spectrum of HIV-1 SUBTYPES HIV genotypes A - D, AE, F, AG-GH, Group N, Group O, BF, H, K, CRF03_AB. For human diagnosis. Testing principle - extraction of HIV RNA). The set will include reagents for all stages: extraction, amplification for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
5 | Multiplex RT PCR Gastro Intestinal-Bacteria(I) Assay | Reaction type - chain polymerization reaction (PCR) with real-time detection. Reagents for at least 25 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection.Detection in stool of Aeromonas spp., Campylobacter spp., Clostridium difficile toxin B, Salmonella spp., Shigella spp./EIEC, Vibrio spp., Yersinia enterocolitica.EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 2
6 | Multiplex RT PCR Gastro Intestinal Virus Assay | Reaction type - chain polymerization reaction (PCR) with real-time detection. Reagents for at least 25 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection. Detection in stool of Adenovirus, Astrovirus, Norovirus GI, Norovirus GII, Rotavirus, Sapovirus. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 2
7 | RT PCR detection of the tox-bearing gene/nontox-bearing C. diphtheriae/ulcerans | Reaction type - chain polymerization reaction (PCR) with real-time detection, test for qualitative detection of Corynebacterium diphtheriae DNA and genes encoding toxins of Corynebacterium diphtheriae and Corynebacterium ulcerans in the biological material (nasopharyngeal swabs, oropharyngeal swabs, swabs from disease sites, germ culture), using real-time hybridization-fluorescence detection of amplified products, IVD.Reagents for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 2
8 | RT PCR kit detection of Bacillus anthracis | Reaction type - chain polymerization reaction (PCR) with real-time detection. Reagents for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
9 | RT PCR kit detection of Clostridium botulinum human-relevant toxingenes – types A, B, E and F | Reaction type - chain polymerization reaction (PCR) with real-time detection. For food samples. Testing principle performing the amplification reaction with real-time detection. Reagents for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 2
10 | RT PCR Leptospira interrogans Kit | Reaction type - chain polymerization reaction (PCR) with real-time detection. For human diagnosis. Testing principle performing the amplification reaction with real-time detection. Reagents for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 2
11 | RT PCR Yersinia pestis kit | Reaction type - chain polymerization reaction (PCR) with real-time detection. For human diagnosis. Testing principle performing the amplification reaction with real-time detection. Reagents for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier.. The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 2
12 | RT PCR Yersinia enterocolitica kit | Reaction type - chain polymerization reaction (PCR) with real-time detection. For human diagnosis. Testing principle performing the amplification reaction with real-time detection. Reagents for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
13 | RT PCR Brucella abortus kit | Reaction type - chain polymerization reaction (PCR) with real-time detection. Testing principle performing the amplification reaction with real-time detection. Reagents for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
14 | RT PCR Francisella tularensis kit | Reaction type - chain polymerization reaction (PCR) with real-time detection. Testing principle performing the amplification reaction with real-time detection. Reagents for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
15 | RT PCR Legionella pneumophila kit | Reaction type - chain polymerization reaction (PCR) with real-time detection. Testing principle performing the amplification reaction with real-time detection. Reagents for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
16 | RT PCR Coxiella burnetii kit | Reaction type - chain polymerization reaction (PCR) with real-time detection. Testing principle performing the amplification reaction with real-time detection. Reagents for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
17 | RT PCR Rickettsia prowazekii kit | Reaction type - chain polymerization reaction (PCR) with real-time detection. Testing principle performing the amplification reaction with real-time detection. Reagents for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
18 | RT PCR Tick‐borne Encephalitis Virus (TBEV) kit | Reaction type - chain polymerization reaction (PCR) with real-time detection. Testing principle performing the amplification reaction with real-time detection. Reagents for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
19 | RT PCR Bordetella Kit (Bordetella pertussis and Bordetella parapertussis) | Reaction type - chain polymerization reaction (PCR) with real-time detection. Testing principle performing the amplification reaction with real-time detection. Reagents for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
20 | RT PCR Neisseria meningitidis blood kit | Reaction type - chain polymerization reaction (PCR) with real-time detection. Testing principle performing the amplification reaction with real-time detection .For human (blood). Reagents for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
21 | RT PCR Enterovirus kit | Reaction type - chain polymerization reaction (PCR) with real-time detection. Testing principle performing the amplification reaction with real-time detection.. Reagents for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 2
22 | RT PCR EHEC PCR Kit Shiga toxin 1 and Shiga toxin 2 and ipaH | Reaction type - Chain polymerization reaction with real time detection. Test principle - total DNA extraction, amplification reaction with detection of the amplified product. For human diagnosis. Reagents for not less than 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection.. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive | kit | 1
23 | RT PCR Malaria Kit | Reaction type - Chain polymerization reaction with real time detection. Test principle - total DNA extraction, amplification reaction with detection of the amplified product. For human diagnosis. Reagents for not less than 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection.. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive | kit | 1
24 | Real Time PCR V. cholerae DNA kit | Qualitative detection of Cholera toxin, reagents for minimum 50 tests, including controls, compatible with Rotor-Gene -6000 amplifier (0.2 ml tubes). The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate, in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
25 | PCR Real Time detection kit of methicillin-sensitive and methicillin-resistant Staphylococcus aureus DNA and methicillin-resistant coagulase-negative Staphylococcus spp. DNA | Qualitative and quantitative detection of methicillin-sensitive and methicillin-resistant Staphylococcus aureus DNA and methicillin-resistant coagulase-negative Staphylococcus spp. DNA in the biological material (oropharyngeal swabs, bronchoalveolar lavage (BAL), sputum, endotracheal aspirate, bronchial washing fluid, urine (first portion pellet), blood, blood plasma, cerebrospinal fluid (CSF), affected organs and tissues aspirates, washes from healthcare equipment and instruments) using real-time hybridization-fluorescence detection, 110 tests/kit, suitable for different real-time PCR instruments | kit | 1
26 | PCR Real Time detection kit is intended to detect genes of resistance to glycopeptide and beta-lactam antibiotics: van A \ B (vancomycin, teicoplanin); mec A (methicillin, oxacillin); tem, ctx-M-1, shv (penicillins and cephalosporins); oxa-40-like, oxa-48-like, oxa-23-like, oxa-51-like, imp, kpc, ges, ndm, vim (carbapenem). | Reaction type - Chain polymerization reaction with real time detection. Test principle - total DNA extraction, amplification reaction with detection of the amplified product. For human diagnosis. Reagents for not less than 24 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection.. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
27 | PCR Real Time detection kit is intended to detect genes of resistance to glycopeptide and beta-lactam antibiotics: van A \ B (vancomycin, teicoplanin); mec A (methicillin, oxacillin). | Reaction type - Chain polymerization reaction with real time detection. Test principle - total DNA extraction, amplification reaction with detection of the amplified product. For human diagnosis. Reagents for not less than 48 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection.. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 1
28 | PCR test-system determination of R-C1-MDR MBL resistance genes in Enterobacteriaceae in real time | Reaction type - Chain polymerization reaction with real time detection. Test principle - total DNA extraction, amplification reaction with detection of the amplified product. For human diagnosis. Reagents for the determination of resistance genes for R-C1-MDR MBL culture of Enterobacteriaceae, for not less than 45 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection.. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. . | kit | 2
29 | PCR test system determination of KPC / OXA-48 R-C2-MDR resistance genes in real time Enterobacteriaceae | Reaction type - Chain polymerization reaction with real time detection. Test principle - total DNA extraction, amplification reaction with detection of the amplified product. For human diagnosis. Reagents for the determination of resistance genes for KPC / OXA-48 R-C2-MDR resistance genes in the bacterial culture of Enterobacteriaceae, for not less than 45 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection.. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 2
30 | PCR test system determination of resistance genes OXA-23, OXA-58, OXA-40 in Acinetobacter baumanii | Reaction type - Chain polymerization reaction with real time detection. Test principle - total DNA extraction, amplification reaction with detection of the amplified product. For human diagnosis. Reagents for the determination of resistance genes OXA-23, OXA-58, OXA-40in the bacterial culture of Acinetobacter baumanii, for not less than 45 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection.. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | kit | 2
Total PCR items
Rapid Diagnostic Tests (RDTs)
1 | Chlamydia trachomatis Ag RDT | 20 tests/kit | kit | 2
2 | N. gonorrhea (ag) RDT | 20 tests/kit | kit | 2
3 | S. pyogenes (Ag) Rapid-Strep test | 20 tests/kit | kit | 10
Total RDTs
TOTAL LOT: 2 Reagents for diagnostic procedures
LOT: 3 Consumables
1 | Sterile loop from polystyrene, 1 mkl/Anză sterilă din polistirol, 1 µl, | Inoculation loop, 1 µl, for single-use, sterile, | pcs | 60500
2 | Sterile loop from polystyrene, 10 mklAnză sterilă din polistirol, 10 µl, | Inoculation loop, 10 µl, for single-use, sterile, | pcs | 64700
3 | Pipette, pasteur; single-use | Transfer pipette 1 ml, with graduation, sterile, 1pc./paper/plastic packaging, | set | 3000
4 | Autoclavable waste bag | Autoclavable waste bag, PP, the bag must not be completely closed, volume 1.5 L, thickness 50 μm, PP-film for the disposal of contaminated labware, BIOHAZARD sign, with indicator patch field turns into dark, 100/set | pcs | 2000
5 | Autoclavable waste bag | Autoclavable waste bag, transparent, volume 6 L, thickness 50 μm, PP-film for the disposal of contaminated labware, BIOHAZARD sign, with indicator patch field turns into dark ,when autoclaved,autoclavable at 134 °C, 20 min, for autoclaving, the bag must not be completely closed, sterilization indicator patch | pcs | 2000
6 | Autoclavable waste bag | Autoclavable waste bag, transparent, volume aprox. 75 l, thickness 50 μm, PP-film for the disposal of contaminated labware, BIOHAZARD sign, with indicator patch field turns into dark when autoclaved, autoclavable at 134 °C, 20 min, for autoclaving, the bag must not be completely closed | pcs | 10000
7 | Plastic Medical Waste Disposal Bin Box Sharps Container | Plastic Medical Waste Disposal Bin Box Sharps Container, 1 L, BIOHAZARD sign, yellow color, material PP | pcs | 10
8 | Plastic Medical Waste Disposal Bin Box Sharps Container | Plastic Medical Waste Disposal Bin Box Sharps Container, 0.5 L, BIOHAZARD sign, material PP | pcs | 3
9 | Plastic Medical Waste Disposal Bin Box Sharps Container | Plastic Medical Waste Disposal Bin Box Sharps Container, 10 L, BIOHAZARD sign | pcs | 5
10 | Autoclave tape (adhesive tape) | Indicator Tape to seal sterilization packs and identify that adequate temperature has been reached during the sterilization cycle, 19MM Wide x 50 Meter Long Roll | pcs | 10
11 | Sample transport box | Silicone water-tight seal and three lid clasps provide secure closure Carry handle folds neatly into body space allowing the units to be stacked Lid opens a full 180°, allowing total access to contents and making it easier to clean Will accommodate 13 and 16 mm tubes in a 72-place OneRack, or similar sized/ brand of rack Use the separators provided to create compartments, or use empty to transport larger products and equipment, Autoclavable, 38.6 x 20.8 x 17cm | pcs | 3
12 | Biohazard Sign | Caution Biohazard Sign - Yellow & Black, waterproof sticker - 12,7 cm length | pcs | 100
13 | Sharpsafe 2l - Int'l Label (Container /deseuri biologice masa plastica (2l) | Useful / total volume 1.70 L / 2.23 L; Dimensions H x L x D (mm) 177 x 197 x 126; Dimensions of large opening (mm) 55 x 50; Assembly: 4 clicks / Single unit 4 clicks; Double closure (temporary and permanent); Safety flaps; Vertical mounting (wall and trolley); Medical rail mounting (wall and trolley); SafeStand horizontal mounting; SafeTray medical tray. | pcs | 25
Total LOT: 3 Consumables