## Sheet1
WHO ITB/RFQ 2022/EURO/MDA/0009 Laboratory equipment for the NAPH National Laboratories
Annex 1. List of requirements and specifications
# | Item | Technical specification (requirements) | Unit | Quantity
1 | Class II Biological Safety Cabinet (BSC) with support stand | Class IIA2 NSF 49 or class II EN 12469; inward airflow ≥0.40 m/s, external height ≤2200 mm including support stand, available space from the top of the BSC to the ceiling > 400 mm; Internal working area approx. (WxDxH) 1450 mm × 630 mm × 650–750 mm (may differ); Inside finish: stainless steel, type 304, with 4B finish; High optical transmission, UV lamp with light timer, absorption of UV light; minimal reflection; Noise pressure level: ≤60 dB; Internal light: Flicker-free, low-glare, warm-colored light, min. 800 lux, control display, LED light indicator, electronic fan control, flow meter for air inflow velocity, flow indicator or meter for air downflow velocity; Operating hours indicator (counter), pre-filter efficiency: min. 85% for particles of minimum 0.5 microns, class G4. HEPA filter (exhaust air filter); classification at least H14 (EN 1822) metal framed; Air downflow velocity: NSF 49–2002, influx air velocity: the average airflow velocity at front aperture should be 0.51 m/s for class A2 (requires compliance with the manufacturer’s set points, or downflow velocity with a deviation of <0.025 m/s from a nominal set point; Factory Test Report as per EN 12469; airflow velocity between 0.25 and 0.50 m/s; air circulation volume flow: approx.: 1000–1500 m³/h; average airflow velocity at front aperture should be at least 0.4 m/s; assigned protection factor (Apf) 1.5*105; volume of fresh airflow inward 400–700 m³/h; alarms, visible and/or audible, for any unsafe condition of the BSC (e.g. airflow, window position, hardware or software errors); possibility to shut down for cleaning and maintenance; Installation, testing procedures, necessary trainings and maintenance within the guarantee period shall be included to the price and provided by a qualified and certified personnel; User manual should be provided on Romanian or Russian languages. | pcs | 8
2 | Hot Air Sterilization cabinet with forced air convection | Volume: aprox 111 litres; working temperature up to 250°C; Interior: stainless steel, trays/shelves - max. 4, Number of external metal door – 1; inner space approx. (W/D/H) - 540/370/530 Intuitive control; smart process control; multi-lingual communication, acoustic and visual alarm; LED indicator of device functionality, LCD display; durable touch-sensitive foil keyboard; keypad lock to prevent unauthorized access or accidental key presses; Interfaces: USB, Ethernet and USB Host; Installation, testing procedures, necessary trainings and maintenance within the guarantee period shall be included to the price and provided by a qualified and certified personnel; User manual should be provided on Romanian or Russian languages. | pcs | 5
3 | Thermocycler PCR (DNA Amplifier) | Sample capacity: 96 well 0.2 m,l block: 10–100 μl; Thermal block: with 6 (six) or more temperature block zones; LED; Optical detection: at least 6 decoupled filters; Excitation / detection: 450–680 nm / 500–730 nm 96 wells; Data acquisition: full board images; Multiplexing: up to 6 targets; Single-plate analysis: presence or absence, absolute and relative analysis of gene expression, SNP genotyping, and high-resolution melting; Maximum ramp rate: 6.0 °C / sec; Average ramp rate for the sample: at least 3.5 °C / sec Temperature accuracy: 0.25° C or better; Compatible fluoro-forms: FAM / SYBR Green, VIC / JOE / HEX / TET, ABY / NED / TAMRA / Cy3, JUN, ROX /; Texas Red, Mustang Purple ™, Cy®5 / LIZ ™, Cy®5.5; Ability to run and edit directly from the screen and computer; Ability to use portable devices to quickly analyze data; Ability to share data sets and protocols online; Installation, testing procedures, necessary trainings and maintenance within the guarantee period shall be included to the price and provided by a qualified and certified personnel; User manual should be provided on Romanian or Russian languages. | pcs | 3
4 | UV recirculator | UV radiation source: 2 lamps - 25W; Bactericidal UV-C; TUV 25W 1SL/25; UV radiation level 36 mW/cm2/s; Air-flow productivity 29 m3/hour; Full user protection from direct UV light; Lamp service life 9000 h; Timer 1 min–24 hrs non–stop; UV lamp lifetime counter. | pcs | 4
5 | Digital camera | Body type - compact, Waterproof (must support pressure water for disinfection and washing); Image minimum ratio types 1:1, 4:3, 3:2, 16:9; Effective pixels min 15 megapixels; Sensor photo detectors min 16 megapixels; Sensor size: Not smaller than 1/2.3" (6.17 x 4.55 mm); Sensor type: BSI-CMOS; ISO Auto, 100-3200, 6400 available at lower resolutions; Image stabilization: Sensor-shift ; Focal length (equiv.): min 28–140 mm; Optical zoom minimum 5×; Maximum aperture: F3.9–4.9; Autofocus, minimum scenes: Contrast Detect (sensor); multi-area; center; tracking; single; continuous; face Detection; live View; Digital zoom: Not less than (2X); Normal focus range: from 65 cm; Macro focus range: 10 cm and lower; Screen / viewfinder; Articulated Fixed LCD; Screen size: Not less than 3″; Screen, dots: minimum 920,000; Live view; Minimum shutter speed 4 sec; Maximum shutter speed: 1/2000 sec; Continuous shoots: Not less 10 fps; Subject / scene modes; Built-in flash; flash range: Minimum 4.0 m (with Auto ISO); Flash modes: Auto, flash on, flash off, slow synchro; Continuous drive: Not less 10.0 fps; Self-timer; Metering modes: Multi; Exposure compensation ±2 (at 1/3 EV steps); Video: min 1920 x 1080 (60p, 30p), 1280 x 720 (60p), 640 x 480 (30p); format MPEG-4, minimum H.264; Microphone, speakers; Storage types: SD/SDHC/SDXC, Internal; Minimum USB 2.0; HDMI (micro-HDMI); Wireless; Built-In; Remote control (via smartphone); Waterproof to 10m, shockproof min 1.5m, freezeproof not lower than -10C, dustproof; Battery Pack; lithium-ion battery & charger; Weight (Including batteries): Maximum 215 g; Orientation sensor. | pcs | 1
6 | Timer | Count down timer and count up timer (stopwatch) with four figure LCD displays; Up to max. 99 minutes 59 seconds, counts in one second steps; Interruption possible; Acoustic alarm for count-down to zero; Set starting time retrievable; Magnet on the back for simple fixing to metal surfaces; Powered by a battery. | pcs | 10
7 | Thermologger | Measures temperature and relative humidity; measurement Range (T) -18 to +55°C; connection using USB cable; Logging Rate: 10 seconds to 24 hours; start modes - immediate; push-to-start; delayed start; parameter triggered; memory capacity: 1,000,000 total readings; sounder - Integrated alarm sound. | pcs | 10