## Sheet1
WHO ITB 2022/EURO/MDA/0009 Laboratory equipment for the ANSP National Laboratories
Annex 4. Bidder's offer and price schedule.
Please fill in the table below (white columns) and send it back in both in Excell format and signed and stamped in pdf
# | Item | Technical specification (minimum requirements) | Technical specification (supplier proposal) | Unit | Quantity | Unit price, USD VAT exclusive | Amount, USD VAT exclusive | Warranty, months | Term of Delivery
1 | Class II Biological Safety Cabinet (BSC) with support stand | Class IIA2 NSF 49 or class II EN 12469; inward airflow ≥0.40 m/s, external height ≤2200 mm including support stand, available space from the top of the BSC to the ceiling > 400 mm; Internal working area approx. (WxDxH) 1450 mm × 630 mm × 650–750 mm (may differ); Inside finish: stainless steel, type 304, with 4B finish; High optical transmission, UV lamp with light timer, absorption of UV light; minimal reflection; Noise pressure level: ≤60 dB; Internal light: Flicker-free, low-glare, warm-colored light, min. 800 lux, control display, LED light indicator, electronic fan control, flow meter for air inflow velocity, flow indicator or meter for air downflow velocity; Operating hours indicator (counter), pre-filter efficiency: min. 85% for particles of minimum 0.5 microns, class G4. HEPA filter (exhaust air filter); classification at least H14 (EN 1822) metal framed; Air downflow velocity: NSF 49–2002, influx air velocity: the average airflow velocity at front aperture should be 0.51 m/s for class A2 (requires compliance with the manufacturer’s set points, or downflow velocity with a deviation of <0.025 m/s from a nominal set point; Factory Test Report as per EN 12469; airflow velocity between 0.25 and 0.50 m/s; air circulation volume flow: approx.: 1000–1500 m³/h; average airflow velocity at front aperture should be at least 0.4 m/s; assigned protection factor (Apf) 1.5*105; volume of fresh airflow inward 400–700 m³/h; alarms, visible and/or audible, for any unsafe condition of the BSC (e.g. airflow, window position, hardware or software errors); possibility to shut down for cleaning and maintenance; Installation, testing procedures, necessary trainings and maintenance within the guarantee period shall be included to the price and provided by a qualified and certified personnel; User manual should be provided on Romanian or Russian languages. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed | pcs | 8 | Please provide price exl. VAT | #VALUE! | Please provide warranty time in months | Please provide term of delivery in days
2 | Hot Air Sterilization cabinet with forced air convection | Volume: aprox 111 litres; working temperature up to 250°C; Interior: stainless steel, trays/shelves - max. 4, Number of external metal door – 1; inner space approx. (W/D/H) - 540/370/530 Intuitive control; smart process control; multi-lingual communication, acoustic and visual alarm; LED indicator of device functionality, LCD display; durable touch-sensitive foil keyboard; keypad lock to prevent unauthorized access or accidental key presses; Interfaces: USB, Ethernet and USB Host; Installation, testing procedures, necessary trainings and maintenance within the guarantee period shall be included to the price and provided by a qualified and certified personnel; User manual should be provided on Romanian or Russian languages. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed | pcs | 5 | Please provide price exl. VAT | #VALUE! | Please provide warranty time in months | Please provide term of delivery in days
3 | Thermocycler PCR (DNA Amplifier) | Sample capacity: 96 well 0.2 m,l block: 10–100 μl; Thermal block: with 6 (six) or more temperature block zones; LED; Optical detection: at least 6 decoupled filters; Excitation / detection: 450–680 nm / 500–730 nm 96 wells; Data acquisition: full board images; Multiplexing: up to 6 targets; Single-plate analysis: presence or absence, absolute and relative analysis of gene expression, SNP genotyping, and high-resolution melting; Maximum ramp rate: 6.0 °C / sec; Average ramp rate for the sample: at least 3.5 °C / sec Temperature accuracy: 0.25° C or better; Compatible fluoro-forms: FAM / SYBR Green, VIC / JOE / HEX / TET, ABY / NED / TAMRA / Cy3, JUN, ROX /; Texas Red, Mustang Purple ™, Cy®5 / LIZ ™, Cy®5.5; Ability to run and edit directly from the screen and computer; Ability to use portable devices to quickly analyze data; Ability to share data sets and protocols online; Installation, testing procedures, necessary trainings and maintenance within the guarantee period shall be included to the price and provided by a qualified and certified personnel; User manual should be provided on Romanian or Russian languages. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed | pcs | 3 | Please provide price exl. VAT | #VALUE! | Please provide warranty time in months | Please provide term of delivery in days
4 | UV recirculator | UV radiation source: 2 lamps - 25W; Bactericidal UV-C; TUV 25W 1SL/25; UV radiation level 36 mW/cm2/s; Air-flow productivity 29 m3/hour; Full user protection from direct UV light; Lamp service life 9000 h; Timer 1 min–24 hrs non–stop; UV lamp lifetime counter. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed | pcs | 4 | Please provide price exl. VAT | #VALUE! | Please provide warranty time in months | Please provide term of delivery in days
5 | Digital camera | Body type - compact, Waterproof (must support pressure water for disinfection and washing); Image minimum ratio types 1:1, 4:3, 3:2, 16:9; Effective pixels min 15 megapixels; Sensor photo detectors min 16 megapixels; Sensor size: Not smaller than 1/2.3" (6.17 x 4.55 mm); Sensor type: BSI-CMOS; ISO Auto, 100-3200, 6400 available at lower resolutions; Image stabilization: Sensor-shift ; Focal length (equiv.): min 28–140 mm; Optical zoom minimum 5×; Maximum aperture: F3.9–4.9; Autofocus, minimum scenes: Contrast Detect (sensor); multi-area; center; tracking; single; continuous; face Detection; live View; Digital zoom: Not less than (2X); Normal focus range: from 65 cm; Macro focus range: 10 cm and lower; Screen / viewfinder; Articulated Fixed LCD; Screen size: Not less than 3″; Screen, dots: minimum 920,000; Live view; Minimum shutter speed 4 sec; Maximum shutter speed: 1/2000 sec; Continuous shoots: Not less 10 fps; Subject / scene modes; Built-in flash; flash range: Minimum 4.0 m (with Auto ISO); Flash modes: Auto, flash on, flash off, slow synchro; Continuous drive: Not less 10.0 fps; Self-timer; Metering modes: Multi; Exposure compensation ±2 (at 1/3 EV steps); Video: min 1920 x 1080 (60p, 30p), 1280 x 720 (60p), 640 x 480 (30p); format MPEG-4, minimum H.264; Microphone, speakers; Storage types: SD/SDHC/SDXC, Internal; Minimum USB 2.0; HDMI (micro-HDMI); Wireless; Built-In; Remote control (via smartphone); Waterproof to 10m, shockproof min 1.5m, freezeproof not lower than -10C, dustproof; Battery Pack; lithium-ion battery & charger; Weight (Including batteries): Maximum 215 g; Orientation sensor. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed | pcs | 1 | Please provide price exl. VAT | #VALUE! | Please provide warranty time in months | Please provide term of delivery in days
6 | Timer | Count down timer and count up timer (stopwatch) with four figure LCD displays; Up to max. 99 minutes 59 seconds, counts in one second steps; Interruption possible; Acoustic alarm for count-down to zero; Set starting time retrievable; Magnet on the back for simple fixing to metal surfaces; Powered by a battery. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed | pcs | 10 | Please provide price exl. VAT | #VALUE! | Please provide warranty time in months | Please provide term of delivery in days
7 | Thermologger | Measures temperature and relative humidity; measurement Range (T) -18 to +55°C; connection using USB cable; Logging Rate: 10 seconds to 24 hours; start modes - immediate; push-to-start; delayed start; parameter triggered; memory capacity: 1,000,000 total readings; sounder - Integrated alarm sound. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed | pcs | 10 | Please provide price exl. VAT | #VALUE! | Please provide warranty time in months | Please provide term of delivery in days
TOTAL | [6 empty] | #VALUE!