## Sheet1
Annex 4
ITB 2022/EURO/MDA/0022
EU COVID-19 Vaccination and Vaccine Preparedness and Deployment Programme for the Eastern Partnership. Republic of Moldova.
Procurement of automated external defibrillator (AED ) for primary health care facilities f the Republic of Moldova
Price Schedule Form
# | Item | Technical specification | Technical specification (supplier proposal) | Unit | Quantity | Unit price, USD | Amount, USD VAT exclusive | Latest Delivery date
1 | Automated external defibrillator (AED) | Description: Semi-automatic defibrillator with ECG analysis, automatic loading and operator-activated discharge. Manual action to activate defibrillation - yes Voice guidance in Romanian or Russian - yes Maximum energy sequences ≥ 200 J Automatic power setting by the device in 3 steps - indicates the power sequences that are reset The reprogramming of the energy sequences is given by the person authorized by the manufacturer Automatic detection of adult or child electrode type yes Rectilinear biphasic output waveform or equivalent Display: Optional ECG signal display is not required Present heart rate icons - yes, Error display - yes Power Sequence Number Display - yes, Resuscitation time display - yes Display Chest Compression Quality (CPR Example) - yes, Show informational messages - yes Self-analysis, Analysis time <15 s Capturing the ECG signal by means of devibrillation electrodes - yes Daily, weekly, monthly self-testing - yes The presence of the icon that qualifies whether the device is valid for use or not. Charging time ≤10 s Communication interface for the transmission of defibrillation data at least IR technology Internal memory ECG and CPR memory capacity at least 7 hours Audio recording capacity (optional) with ECG and Audio minimum 20 minutes Presence of guiding images / pictograms showing steps for patching the patient with defibrillation - yes The presence of the headrest that allows the correct arrangement of the head for the mouth-to-mouth breathing procedure - yes Accumulators: Stand-by life of at least 5 years Operating time> 200 discharges or> 10 continuous hours of ECG monitoring Possibility to buy batteries on the local market - yes Weight ≤ 3.5 kg Accessory Disposable adult defibrillation pads, with a shelf life of 5 years at the time of delivery to include the barrier mask (for the mouth-to-mouth procedure) a disposable disposable device, scissors, disposable gloves. ≥ 2 pcs These clear pictograms refer to the correct placement on a pediatric patient, single use, with a shelf life of 2 years at the time of delivery ≥ 1 pc The presence of clear icons with references to the correct placement on the patient Defibrillator transport / storage bag The bag should be provided for transporting and storing the defibrillator. Equipped with pockets for all accessories Device warranty - minimum 5 years | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed | pcs | 70 |  | 0 | 31.12.2022
2 | AED wall cabinet | Cabinets on wall to fit AED Cabinets protect AED while making it highly visible Inscription/Sticker (AED) on the cabinet included Material : powder coated steel/stainless stel/ | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed | pcs | 70 |  | 0 | 31.12.2022
3 | Reserve Pads (5 reserve pads for each AED | Suitable replace Pads for AED | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed | pcs | 350 |  | 0 | 31.12.2022
4 | Reserve pediatric Pads (5 reserve pads for each AED | Paediatric electrodes designed for childrens under the age of min 7 years | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed | pcs | 350 |  | 0 | 31.12.2022
5 | AED carry bag | Bag suitable for AED to be included | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed | pcs | 70 |  | 0 | 31.12.2022
6 | Shipping and handling | Transport cost shall be included in the cost |  | lumpsum |  |  | 0
Total | [6 empty] | 0 |  | 0
We declare that all the information and statements made in this Proposal are true and we accept that any misinterpretation or misrepresentation contained in this Proposal may lead to our disqualification .
We offer to provide goods in conformity with the Bidding documents, and in accordance with the Terms of Reference
Our Proposal shall be valid and remain binding upon us for the period of time specified in the Bid requirements.
I, the undersigned, certify that I am duly authorized by WHO ITB 2022/EURO/MDA/00022 to sign this Proposal and bind it should WHO accept this Proposal.
Name: _____________________________________________________________
Title: _____________________________________________________________
Date: _____________________________________________________________