## WHO_MDA_LOT 1_ICU bed_techspecs
Annex 1
ITB 2023/EURO/MDA/0016
LOT 1
Procurement of the Intensive Care Unit patient beds for the Causeni Hospital and Ignatenco Municipal Hospital
List of requirements and specifications
ICU bed for adults and pediatric patients
NAME, CATEGORY AND CODING
| Parameters |  |  | Minimum expected Specification | Proposed Specificacion (to be completed by the bidder) | Reference document / brochure / page where provided information may be verified by the evaluation committee
| 1 | Generic name | Critical Care Bed for adults and pediatric patients
PURPOSE OF USE
| 2 | Clinical or other purpose | Hospital bed to be used in Critical Care
| 3 | Level of use | District hospital, provincial hospital, national hospital, specialized hospital.
| 4 | Clinical department/ward | Intensive Care Unit, Emergency, Recovery
TECHNICAL CHARACTERISTICS
| 3 | Platform | minimum 4 section | required
|  |  | every section can be removed for cleaning | required
| 4 | Bed motion | Fully motorized electric bed | required
|  |  | backrest movement | required
|  |  | legrest movement | required
|  |  | bidirectional tilt movement | required
|  |  | up and down movement height adjustment | required
| 5 | Minimum bed positions | Tredelenburg | required
|  |  | Reverse tredelenburg | required
|  |  | Safety CPR release button/lever | required
|  |  | Vascular | required
|  |  | Fowler | required
|  |  | Chair/Cardiac | required
| 6 | Bed position controled by: | minimum 1 integrated patient control panel in side rail | required
|  |  | minumium 1 patient wired control panel | required
|  |  | minimum 1 nursing control panel on each side of the bed | required
|  |  | foot pedal (up and down movement) | required
| 7 | Side rails | minimum 4 side rails | required
|  |  | every side rail can be moved separately minimum up and down position | required
|  |  | patient support grips minimum one on each proximal compartment side rails | required
| 8 | Nurse control | Patient control lockouts | required
| 9 | Bed-ends | Removable headboard and footboard | required
|  |  | protective bumpers for all corners | required
|  |  | sockets to support IV poles minimum 2 | required
| 10 | Scale | integrated | 1
| 11 | Night light functionality | integrated | 1
| 12 | X-ray tray | integrated | 1
| 13 | Linen holder | integrated | 1
| 14 | Castors | minimum 4 swivel castors | required
|  |  | minimum 2 of which with automatic brakes, total lock type (weel and rotation lock) | required
| 15 | Alarms | Bed-exit, power failure, low battery" | required
|  |  | Fifth wheel | required
| 16 | IV Pole | >2 | required
| 17 | Angle indicators | minimum 1 digital (display) | required
|  |  | minimum 4 mechanical | required
| 18 | Materials: | Frame: epoxy coated tubular steel. | required
|  |  | High resistance to corrosion | required
|  |  | Cover: plastic, flexible highly tear resistant, anti-static, flame retardant, disinfectant- and liquid proof, washable | required
PHYSICAL CHARACTERISTICS
| 19 | Dimensions: | Sleeping surface minimum dimensions: 210x100 cm (l x w) | required
|  |  | low position floor to platform maximum 45cm | required
| 20 | Maximum patient weight | minimum 180Kg | required
UTILITY REQUIREMENTS
| 21 | Electrical | Power: 230 VAC | required
|  |  | Battery Backup | integrated
|  |  | Power supply cord | included
|  |  | Type of protection against electrical current | type 1
|  |  | Protection against water and dust | IP X4
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
| 22 | Accessories/ mattress | included | required
|  |  | minimum height 15 cm | required
|  |  | high-density viscoelastic polyurethan foam, density 27-33 kg/m3 | required
|  |  | Mattress cover removable via side zipper | required
|  |  | Mattress dimension fully fits the platform dimesions | required
TRAINING, INSTALLATION AND UTILISATION
| 23 | Transportation | Supplier must include the transportation to the final healthcare facility | required
| 24 | Installation | Supplier must perform the installation, safety and operation checks before handover | required
WARRANTY AND MAINTENANCE
| 25 | Warranty | At least 24 months | required
DOCUMENTATION
| 26 | Documentation requirements | All supporting documentation, operation, service and user manuals must be presented at least in English and, when available, in the official language of the country in which the equipment will be used, including: | required
SAFETY AND STANDARDS
| 27 | Regulatory Approval / Certification | Free Sales Certificate (FSC) provided by any of the following countries: Australia, Canada, Japan, USA and/or European Community (e.g. CE certificate given by a third certified party for the specific medical devices proposed (not only declaration of conformity)). | required
| 28 | Standards for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485:2016 Medical devices – Quality management systems – Requirements for regulatory purposes) | required

## WHO_MDA_LOT 2_C Arm_techspecs
Annex 1
ITB 2023/EURO/MDA/0016
LOT 2
Procurement of the C-arm radiodiagnostic system for the Causeni Hospital
List of requirements and specifications
C-arm radiodiagnostic system
NAME, CATEGORY AND CODING
| Parameters |  |  | Minimum expected Specification | Proposed Specificacion (to be completed by the bidder) | Reference document / brochure / page where provided information may be verified by the evaluation committee
| 1 | Generic name | C-arm radiodiagnostic system
PURPOSE OF USE
| 2 | Clinical or other purpose | C-arm radiodiagnostic system, intended for performing radiodiagnostic examinations during surgical interventions.
| 3 | Level of use | District hospital, provincial hospital, national hospital, specialized hospital.
| 4 | Clinical department/ward | Intensive Care Unit, Emergency, Recovery, Medical Imaging, Outpatient Services, Inpatient Care, Specialized Diagnostics, Specialized Treatment
TECHNICAL CHARACTERISTICS
| 5 | X-ray generator | high frequency | 60-100 KHz
|  |  | KV range | 40-120
|  |  | generator power | ≥20 kW
|  |  | cooling system | controlled by the processor
| 6 | Detector technology | Image intensifier | required
| 7 | Image intensifier | triple boost mode | minimum 9" HRC
|  |  | Constant filtering | ≥ 0.1 mm Cu + 2.1 mm Al
| 8 | X-ray tube | type | with rotating anode
| 9 | Working regimes | Fluoroscopy | Normal
Pulsed fluoroscopy of 1-30 pulses per second
|  |  | Radiographic | Yes
|  |  | Digital Spot | Yes, Focal Sports 0.3 to 0.8 mm
|  |  | Reduced radiation dose | Yes
| 10 | Control | Automatic Exposure Control (AEC) | yes
| 11 | The C spring | Minimal patient space | ≥ 70 cm
|  |  | the minimum depth | ≥ 70 cm
|  |  | Horizontal and vertical tube movement | hor.>200cm & Vertical tube movement: >60cm"
|  |  | The minimum angle of inclination | ≥±90°
|  |  | Minimum orbital rotation angle | ≥±180°
|  |  | Minimum positional angulation angle | ≥±10°
|  |  | Electric arm movement | Yes
|  |  | lockable wheel stand | Yes
| 12 | Control panel | Touchscreen technology | Yes
| 13 | Remote | Remote control/pedal | Yes
| 14 | Digital camera | built-in CCD type | minimum resolution 1024x1024
|  |  | Image depth, bit | Minimum 16 bits
|  |  | Image processing time on the display | No more than 5 sec
| 15 | Views | booth with at least 2 monitors (or equivalent) | ≥19”
|  |  | Touchscreen type monitor | Yes
| 16 | Software | Compatible with hardware | included
| 17 | Image System & Connectivity | Image storage | Minimum 50,000 full resolution images
|  |  | Processing imported images | DICOM 3.0 compatible
|  |  | Export images | USB
|  |  | DAP meter | Yes
UTILITY REQUIREMENTS
| 18 | Electrical | The voltage of the electric current | 230 V ±10%; 50 Hz
|  |  | Plug type in compliance with required local/national standards and regulations | required
|  |  | Appropriate external device to protect the equipment against over-voltage and over-current line conditions (between plug and socket) | required
TRAINING, INSTALLATION AND UTILISATION
| 19 | Transportation | Supplier must include the transportation to the final healthcare facility | required
| 20 | Installation | Supplier must perform the installation, safety and operation checks before handover | required
| 21 | Training of user/s | Training of users in hardware and software (including applications): | required
|  |  | • Basic operation of the system | required
|  |  | • Advanced operation of the system | required
|  |  | • Maintenance routine | required
WARRANTY AND MAINTENANCE
| 22 | Warranty | At least 24 months | required
| 23 | Preventive maintenance tasks | Included in the warranty twice a year | required
| 24 | Software / Hardware upgrade availability | (any avaiable updated during the warranty period) Included | required
DOCUMENTATION
| 25 | Documentation requirements | All supporting documentation, operation, service and user manuals must be presented at least in English and, when available, in the official language of the country in which the equipment will be used, including: | required
|  |  | • Manual procedures required for local calibration and routine maintenance | required
|  |  | • Troubleshooting protocols | required
|  |  | • Cleaning and disinfection protocols | required
|  |  | • List of spares and accessories, with their part numbers and cost. | required
SAFETY AND STANDARDS
| 26 | Regulatory Approval / Certification | Free Sales Certificate (FSC) provided by any of the following countries: Australia, Canada, Japan, USA and/or European Community (e.g. CE certificate given by a third certified party for the specific medical devices proposed (not only declaration of conformity)). | required
| 27 | Standards for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485:2016 Medical devices – Quality management systems – Requirements for regulatory purposes) | required