## List1
Annex 1
ITB 2023/EURO/MDA/0020
LOT 1
Procurement of the Surgical Equipment for the Insitute of Oncology
List of requirements and specifications
Infusion pump with patient-controlled analgesia (PCA) mode
| Description | Advanced feature syringe infuser with Patient Controlled Analgesia (PCA) mode that allows intravenous administration of medical fluids with high accuracy.
| Parameter |  |  | Specification
| General characteristics | Type |  | with one channel
|  | Display |  | LCD
Touch screen
the possibility to adjust brightness
|  | Possibility to place and fixate on a column |  | Yes
| Display | Data displayed | Infusion rate | Yes
|  |  | Infusion volume | Yes
|  |  | Infusion time | Yes
|  |  | Time till end of infusion | Yes
|  |  | Syringe size | Yes
|  |  | Syringe manufacturer | Yes
|  |  | The level of occlusion | Yes
|  |  | KVO rate | Yes
|  |  | Date, time | Yes
|  |  | Battery charge level | Yes
|  |  | Events | Yes
|  |  | Occlusion pressure | Yes
|  |  | Alarms | Yes
| Pump properties | Flow rate | Range | 0.01 - 2000 ml/h
|  |  | The increment step | 0.01 mL/h
|  | Bolus dosing | incrementing | 0.1 - 2000 ml/h
|  | KVO rate |  | 0.01 - 3 ml/h, with increments of 0.01 ml/h
|  | Accuracy |  | ≤2%
|  | Compatible with all required syringe sizes produced by different manufacturers |  | min. 8 producers, with the presentation of the list of producers
|  | Ability to add new types of syringes to the library |  | Yes
|  | Syringes accepted |  | 2, 3, 5, 10, 20, 30, 50, 60
|  | Automatic syringe detection |  | Yes
|  | Maximum occlusion pressure level |  | 50 - 1000 mmHg
min. 15 levels
|  | Possibility of adding other drug names to the library |  | Yes
|  | Protection against liquids entering the device |  | minimum IP33
| Modes of infusion | Perfusion rate |  | Yes
|  | Dosing over time |  | Yes
|  | Ascending and descending mode |  | Yes
|  | Intermittent mode |  | Yes
|  | Micro-infusion dosing |  | Yes
|  | Patient controlled analgesia (PCA) |  | Yes
| Alarms and indicators | Occlusion |  | Yes
|  | System error |  | Yes
|  | Empty syringe |  | Yes
|  | Infusion soon to end |  | Yes
|  | Discharged battery |  | Yes
| Audible alarm | Adjustable volume |  | Yes
|  | Silent mode button |  | Yes
| Log book | Number of events stored |  | Yes
|  | Stored events | Buttons pressed | Yes
|  |  | Error codes | Yes
|  |  | alarms | Yes
|  |  | Injected drugs | Yes
|  |  | Amount infused | Yes
|  |  | settings | Yes
|  |  | Bolus requested | Yes
| Power | Electrical network | 220 V, 50 Hz | Yes
|  | Internal battery | Rechargeable | Yes
|  |  | Operating time | ≥ 5 h at 5 mL/h
Accessories included separately for each pump purchased
| Support | system for fixing the pump to the infusion stand |  | Yes
| Controller for patient-controlled analgesia module |  |  | Yes
| Manufacturer's certificate attesting product quality ISO 13485 - valid - copy confirmed by the participant's signature and stamp. |  |  | required
| CE certificate or CE declaration of conformity depending on the assessment of conformity with the corresponding annexes for the products offered - valid - certified copy signed and stamped by the participant. |  |  | required
| User manual, service manual, according to law 102 of 2017 with all amendments, Installation, commissioning, training of users and technicians by the winning participant - mandatory |  |  | required
| Minimum 24-month warranty from the moment of commissioning. |  |  | required