## List1
Annex 1
ITB 2023/EURO/MDA/0020
LOT 2
Procurement of the Laboratory Equipment for the Insitute of Oncology
List of requirements and specifications
Plasma sterilizer
Plasma sterilizers are used for rapid cold sterilization in hospitals, laboratories and dental clinics for all heat and moisture sensitive equipment that is resistant to H2O2
Parameter | Specification
Loading | horizontal
Interior construction | stainless steel
Exterior Construction | galvanized metal or stainless steel
Configuring | one door
Alarms | acoustic, visual
Sterilization time adjustable | Yes
Preset sterilization regimes | ≥ 2 programs
Sterilization at low temperatures | Yes
Electronic control | Microprocessor
Sterilizing agent | H2O2
Thermal printer | Yes
Process monitoring | on LCD screen or LED display
Inner volume | from 80 to 100 liters
Support with compatible shelves for the placement of instruments and cases | 1 supports
shelves min. 2 pcs.
Sterilization time | Short cycle max. 35 minutes
Standard cycle max. 60 minutes
Temperature range: | 30-60°C
Data logger | Yes
Consumables for 200 sterilization cycles (tests, sterilization agent, etc.) | Yes
User manual, service manual, according to law 102 of 2017 with all amendments, installation, commissioning, training by the winning participant - mandatory training for users - mandatory. | Required
Standards and regulations: 93/42/EEC; EN ISO 13485, ISO 14937 validated sterilization cycles and CE (LVD IEC 61010-2-040, LVD IEC 61010-1) certification | Required
Minimum warranty of 24 months from the moment of commissioning. | Required