## Sheet1
Annex 1
ITB 2023/EURO/MDA/0020
LOT 1
Procurement of the Surgical Equipment for the Insitute of Oncology
List of requirements and specifications
Electrocautery / Electrosurgery unit with accessories
| Parameter | Specification
| Description | A universal electrosurgical platform providing adapted energies in sealing and cutting vessels, intended for laparoscopic and open surgical procedures, based on energy generation for general application, urology, including gynecology, colorectal, cardiac, ENT.
|  | Monopolar with international 3-pin connection |  | 2 outlets
| Output channels | Bipolar and vessel sealing |  | 2 outlets
|  | Hand switch |  | Yes
|  | Bipolar footswitch and vessel sealing |  | Yes
|  | Monopolar footswitch |  | Yes
| Working modes | Monopolar | Pure cut | ≥300W
|  |  | Blend cut | ≥ 200 W
|  |  | Coagulation Lightning | ≥ 120W
|  |  | Coagulation Soft | ≥120 W
|  |  | Hemostasis spray | ≥ 100W
|  | Bipolar | Pure cut (resection or vaporization) | ≥ 370W
|  |  | Coagulation cut (resection or vaporization) | ≥150 W
|  |  | Coagulation | ≥ 90W
|  | Vessel Sealing (Tissue Fusion, Ligature Vessel Sealing, or equivalent) | Maximum power | 350 W
Seals blood and lymphatic vessels up to and including 7 mm in diameter
| Work schedules | Surgical applications |  | minimum 50 programs
| Screen | LCD touchscreen with the ability to control at least 4 different working modes (Monopolar, Bipolar, Sealing vessels, laparoscopic)
Separate display in 4 quadrants
| Self-test function |  |  | Yes
| Indicators |  |  | acoustic
visual
| Sound volume control |  |  | Yes
| Built-in software for bipolar resection in saline environment |  |  | Yes
| Supply voltage |  |  | 220 - 240 V, 50 Hz
Additional functions
The work sockets are of the "SMART" type, they automatically recognize the work tools and have automatic energy configuration
Automatic recognition of completion of the Vessel Sealing cycle (tissue fusion, vessel sealing) with interruption of power supply to the instrument and cycle completion sound
Interactive error communication, including on-screen corrective actions in case of non-completion of the Vessel Sealing cycle
Automatic power adjustment system based on tissue impedance within selected power limits
Accessory
| Power cable |  |  | 2 pcs.
| Mobile cart with wheels for transporting the electrosurgery unit |  |  | 1 pc.
| Single-circuit footswitch for bipolar control, with connection cable included |  |  | 2 pcs.
| 2-circuit footswitch for monopolar/bipolar control with connection cable included |  |  | 2 pcs.
| Reusable neutral electrode, with minimum 3 m cable included |  |  | 3 pcs.
| Reusable monopolar handle (pencil) for the electrode, with at least two operating buttons, min. 3 m cable |  |  | 10 pcs
| Set of electrodes (min 3 pcs.) of different sizes and shapes (ball, loop, blade, needle) |  |  | 5 set
| Monopolar cable of at least 3 m length (for connecting classical instruments) |  |  | 2 pcs.
| Autoclavable bipolar cable with a minimum length of 3 m (for connecting classical instruments) |  |  | 2 pcs.
| Autoclavable bipolar forceps of different lengths between a minimum of 18 cm and a maximum of 26 cm, used in open surgery |  |  | 2 pcs.
| Forceps for cutting and sealing blood vessels (integrated cutting system), used in open surgery, minimum length 20 cm, jaws 5 mm, reusable, autoclavable |  |  | 2 pcs.
| Forceps for cutting and sealing blood vessels (integrated cutting system), used in open surgery, minimum length 20 cm, jaws 10 mm, reusable, autoclavable |  |  | 2 pcs.
| Autoclavable cable for open surgery vascular sealing forceps (if applicable) |  |  | 4 pcs.
Other requirements
| 13485 standard, CE certification, Standards and classifications: meets all the conditions required by the standards IEC 60601-1 edition 2.0 and 3.1, IEC 60601-2-2 edition 4.0 and 5.0, IEC 60601-1-2 edition 2.1 and 3.0 and 4.0, IEC 60601-1-8 edition 2.1. |  |  | required
| The platform offers a high degree of protection against electric shocks and can be used in cardiac procedures CF Equipment/Defibrillator Proof (IEC 60601-1, IEC 60601-2-2, and ANSI/AAMI HF18) |  |  | required
| Electrical safety Class 1, CF type |  |  | required
| User manual, service manual, according to law 102 of 2017 with all amendments., Installation, commissioning mandatory |  |  | required
| Training for users and technicians, performed by a specialist in the field - mandatory. |  |  | required
| Minimum 24 months warranty from the moment of commissioning. |  |  | required