## Sheet1
Annex 1
ITB 2023/EURO/MDA/0020
LOT 2
Procurement of the Laboratory Equipment for the Insitute of Oncology
List of requirements and specifications
Autostainer
| Description | Slide staining device according to immunohistochemical protocols
| Parameter | Specification
| System type | open
| Allowing use with any type of reagent for any of the staining steps | Yes
| Slide capacity | minimum 45 slides
| Reagent capacity | minimum 45 reagents
| Allowing special staining programming | Yes
| Independent programming of each slide | Yes
| Labeling of each slide and reagent container | Yes
| Bar code reader | Yes
| PC workstation | Yes
| Specialized software | Yes
| Label printer | Yes
| Reactive containers | Yes, minimum 200 pcs
| UPS. autonomy minimum 1 hour | Yes
| How to deparaffinize and unmask the antigen before staining | performs simultaneously both deparaffinization and antigen unmasking
Display touch screen
Predefined programs available but also programs created by the user
2 reaction baths for a minimum of 24 slides each
Each bath can be programmed and allow operation either according to the same program or according to different programs
Minimum reagent tank volume: 1.5 L
Each bath must have a maximum load indicator
allows setting the time from 1 min up to minimum 5 hours
adjustable temperature from ≥18 °C to≤200 °C
display resolution +0.1°C below 99.9°C, 1°C above 100°C
RS232 communication interface or equivalent
stainless steel interior/exterior construction
overtemperature safety system
| Installation, commissioning, training by the winning participant - mandatory | Required
| User training upon installation and upon request - mandatory. Preventive maintenance services during the warranty period. | Required
| Quick guide in Romanian laminated attached to the device | Required
| Certificate from the manufacturer certifying product quality ISO 9001; CE, ISO 13485 – valid – copy confirmed by participant's signature and stamp. | Required
| CE certificate or EC declaration of conformity depending on the assessment of conformity with the corresponding annexes for the products offered - valid - copy confirmed by the participant's signature and stamp. | Required
| User manual, service manual, according to law 102 of 2017 with all amendments | Required
| Minimum 24-month warranty from the moment of commissioning. | Required