## Sheet1
Annex 1
ITB 2023/EURO/MDA/0020
LOT 2
Procurement of the Laboratory Equipment for the Insitute of Oncology
List of requirements and specifications
Autoscanning epi fluorescent trinocular microscope for FISH
| Application | Multiplex (multicolor) fluorescence in situ hybridization (FISH) and karyotyping
| Parameter | Specification
| Optical system | Infinite optical correction system
| Table | mechanical for right hand operation with coaxial drive for X and Y movement
| Specimen fixative (slides) | universal for a single specimen
| Revolver head | for 4 objectives
| Objectives IOS Plan Achromat: | 4x/0.1, 10x/0.25, 40x/0.65, 100x/1.25 (Oil Immersion)
| Trinocular tube | with an inclination angle of 30°
rotation, 360°
with the adjustment of the interpupillary distance in the minimum range of 52-75mm
| Eyepiece - wide field | WF 10x/22 – 2 pcs. (at least one with focus)
| Focus system | Coaxial knobs on both macrometric and micrometric sides. System for adjusting the tension and the maximum distance traveled on the Z axis.
| Filters | B (Blue), G (Green), V (Violet), O (Orange), R (Red) and UV
| Lighting for fluorescence | 6 Position LED Fluorescent Attachment, Interchangeable LED Filter Sets
| Power unit | BUILT
| Filters | B (Blue), G (Green), V (Violet), O (Orange), R (Red) and UV
| Digital camera | Digital camera resolution: min 1.4 MP
Monochrome with cooling
Signal output: USB 3.0
Dynamic Range (DB): 50 - 80
Sensor size: 1/2.5 - 2/3”
Sensor technology: CMOS or CCD
Image format: minimum 11/8
Pixel size: minimum 6 μm x 6 μm
Exposure time: minimum 20 μsec to maximum 5 sec
Camera power: PC USB
USB 3.0 cable
Micrometer blade
External power supply
| Software for image analysis | Application: detection, analysis and display of structural and molecular features of multicolor FISH specimens, karyotyping
Improving the original image quality
Licensed software with lifetime license
Database of digital preparations
| Computer | Intel® Core™ i5-5000 processor or equivalent
SSD ≥ 1 TB
RAM ≥ 8 GB
Full HD monitor, ≥ 21”, LCD ≥ 1920x1080
USB ≥ 4 ports, at least one 3.0
Windows 10 operating system or equivalent
Accessories: Mouse, keyboard
| Warranty period not less than 24 months from the moment of installation | Required
| Installation, commissioning, training by the winning participant - mandatory | Required
| Training for users performed by a specialist in the field - mandatory. | Required
| Preventive maintenance services during the warranty period | Required
| Quick guide in Romanian laminated attached to the device | Required
| Certificate from the manufacturer certifying the quality of the ISO 12485 product - valid - copy confirmed by the participant's signature and stamp. | Required
| CE certificate or EC declaration of conformity depending on the assessment of conformity with the corresponding annexes for the products offered - valid - copy confirmed by the participant's signature and stamp. | Required
| User manual, service manual, according to law 102 of 2017 with all amendments, | Required