## LOT 1
Annex 4, LOT 1
ITB 2023/EURO/MDA/0022
Procurement of laboratory consumables and devices, for the National Agency for Public Health
Price Schedule Form
#No | Item description | Technical Specification | Technical specification (supplier proposal) | UM | Quantity | Price/unit USD, VAT exclusive | Amount, USD | Comments
Devices for diagnostic procedures
1 | Water distiller with built-in storage tank/Distilator de apă cu rezervor de depozitare incorporat | Productivity: -12-25 l/hour, Built-in storage tank with drain valve- 24-50 liters, Electrical conductivity- 2.3- 2.5 µS/cm, Power supply- 380-400 V / 50-60 Hz, 18 kW, three-phase, Type of water treatment - single distillation, Electronic control system - automatic stop/restart of the water level switch in the storage tank (empty/full), Electronic water impurity detector with visual alarm, stop thermostatic at low water level, protection against leaks with solenoid valve; Stainless steel heating element, evaporator, storage tank and condenser; Double-walled housing in electrolytically galvanized steel, electrostatically powder coated with epoxy resin; Easily accessible cleaning of the evaporator, tank; Possibility of mounting on the table and on the wall; Complete accessories - power cable, electromagnetic valve with built-in connection, pressure hose 1/2" (inner ø approx. 12.7 mm) for connection to the water network (inlet) and hose with 3/4" connection (ø approx. 19.0 mm) for the cooling water outlet; Spare parts – heating elements (heaters); The warranty period is at least 2 year from the moment the equipment is installed Installation and staff training at the final recipient's premises. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | pcs | 10 | 0 | 0 | please provide earliest delivery time, but not later than 31.08.2023
2 | Water filtration device with 3-place vacuum pump for microbiological water analysis/ Dispozitiv pentru filtrarea apei cu pompă de vid cu 3 locuri pentru analiza microbiologică a apei | Autoclavable stainless steel collector (pump) with 3 places complete with 3 autoclavable stainless steel funnels with a volume of 500 ml: The presence on the inside of the palms signs that allow you to control the filled volume of 100 ml., 250 ml., stainless steel containers for funnels; the support for the filter with a diameter of -47-50 mm which presents a sieve made of sintered metal (to allow sterilization by buckling); Compatibility: filter membranes of Ø47-50 mm; the presence of the filter membrane recovery system; autoclavable/processable silicone connecting tube (hose) with a diameter of 4x14 mm- 10-15 m; Diaphragm vacuum pump 18-20 L/min, vacuum -0.8-1 kgс/см2 (78-98kPa- according to the requirements of ISO 8199), supply source – 220 V, air flow 4.0-6.0 l/min., energy consumption 30 W; "Duran" flask (vessel) for collecting filtered water - 2000-2500 ml - 2 pcs, silicone/rubber stopper for the Duran flask - 2-3 pcs; the presence of a special moisture separator filter (water trap); Rubber plug for Duran flask with d. 65x45mm - 2-3pcs; Hose connectors with a diameter of 100x12 mm; The user manual translated into Romanian. The warranty period is at least 2 year from the moment the equipment is installed Installation and staff training at the final recipient's premises. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | pcs | 5 |  | 0
3 | Electronic laboratory balance with accuracy: 0.01/Balanță electronică de laborator cu precizia: 0,01 | Balance with the precision of 0.01g. Digital setting, for weighing up to 1000g Division: 0.01g, calibration mode - internal automatic. Stabilization time 2-4 sec. Stainless steel weighing platform. Gross/net weighing, zeroing and maintenance (country). working temperature: +10...40ᵒC. Display: Illuminated LCD. Power supply 220 V Calibration certificate from the manufacturer, ISO 9001 and 13485 certificate from the manufacturer, Declaration of conformity. At the time of delivery, the equipment must be included in the "State Register of measuring instruments" of the National Institute of Metrology. , User manual translated into Romanian/Russian. The warranty period is at least 2 year from the moment the equipment is installed Installation and staff training at the final recipient's premises. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | pcs | 4 |  | 0
4 | Thermostat (laboratory incubator) with natural convection/Termostat (incubator de laborator) cu convectie naturala | Characteristics: Internal volume at least 240 liters Digital temperature setting with an accuracy of 0.1oC stainless steel interior Textured stainless steel exterior, Alarm system, Overheat protection, Internal door in heat-resistant glass with magnetic closure, 2 chrome shelves as standard equipment Temperature range (°C) +5...+80 Power supply 50/60Hz(V) 230 Nominal power (W) at least 700 - Protection class - IP 42 - staff training, - the presence of the metrological verification certificate Optional accessories: thermostat support on wheels perforated stainless steel shelves -2 pcs., programmable clock 0-12 hours The warranty period is at least 2 year from the moment the equipment is installed Installation and staff training at the final recipient's premises. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | pcs | 4 |  | 0
5 | Laboratory ionomer/Ionomer de laborator | Range: pH - (-2,000…+19,999) Millivolt units - (-3000-2000) Automatic or manual temperature compensation within the limits of -20… +150 ᵒC. Manual or automatic pH setting with buffer solutions. Automatic calibration. Power supply 220 V Display - graphic LCD Contents of the equipment package: compatible convector, universal tripod for electrodes, electrodes 2 pcs. (selective ion NO3, pH), accessories kit (set of cables for verification and buffer solutions with pH: 1.65; 4.00; 6.86; 9.18) Instructions for use Initial metrological verification bulletins Manual of use translated into Romanian & Russian. Installation and staff training at the user's premises | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | pcs | 4 |  | 0
6 | Direct-Q® 3 UV Water Purification System | Direct-Q® 3 UV Water Purification Systems with built-in UV dual wavelength lamp produce both RO pure (Type 3) and ultrapure (Type 1) water directly from the tap at a flow rate of 3 L/hr. • 185/254 nm UV lamp reduces the levels of organics and bacterial for critical applications • Install under or on the bench or wall-mount • Can be equipped with an integrated (6 L) or external (30 or 60 L) reservoir • Ultrapure water quality can be fine-tuned to match application needs using Application-Pak polishers to remove specific contaminants Ionic rejection by RO ≥ 96% Organic rejection by RO ≥ 99% for MW ≥ 200 Particulate rejection by RO ≥ 99% Product Water DNase < 5 pg/mL (with Biopak® Polisher) Product Water Microorganisms < 0.1 CFU/mL (with Millipak® or Biopak® point-of-use filters) Product Water Pyrogens < 0.001 EU/mL (with Biopak® Polisher) Product Water RNase < 1 pg/mL (with Biopak® Polisher) Physicochemical Information Product Water Conductivity at 25°C 0.056 µS/cm Make-up flow rate 3 L/h Flow Rate 3 L/h Product Water Resistivity at 25 °C 18.2 MΩ·cm Product Water TOC < 5 ppb Voltage 100–230 V / 50–60 Hz The warranty period is at least 2 year from the moment the equipment is installed Installation and staff training at the final recipient's premises. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | pcs | 1 |  | 0
7 | Laboratory centrifuge/Centrifuga de laborator | Laboratory centrifuge with max. 8000 rpm, adjustable. Rotor with min. 4 slots with adapters for tubes with a volume of 5ml, 10ml, 15ml, 50ml, 100ml. Protection mechanism that prevents the centrifugation from starting when the lid is lifted. Protection system for stopping the centrifugation in case of exceeding the maximum speed or in case of an imbalance of the rotor, blocking the cover during operation. Stainless steel inner chamber Initial metrological verification bulletins User manual translated into Romanian & Russian. The warranty period is at least 2 year from the moment the equipment is installed Installation and staff training at the final recipient's premises. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | pcs | 3 |  | 0
8 | Drying cabinet/Dulap de uscare | Stainless steel working chamber Operating temperature range from +50ºC to +300ºC Fast heating Forced air convection (fan) increases drying efficiency and accelerates heating, ensures uniform temperature distribution and economical operation Manual control of the exhaust opening in the rear wall Loading at different levels through removable shelves Controller for digital fan speed setting 0-100% in 10% steps (in basic version) Comfortable swivel handle Covering the furnace body with heat-resistant powder paint Additional options: the viewing window Additional shelves or trays for bulk materials Electric oven certification The warranty period is at least 2 year from the moment the equipment is installed Installation and staff training at the final recipient's premises. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | pcs | 3 |  | 0
| TOTAL LOT 1 Devices for diagnostic procedures | [5 empty] | 0
We declare that all the information and statements made in this Proposal are true and we accept that any misinterpretation or misrepresentation contained in this Proposal may lead to our disqualification .
We offer to provide services in conformity with the Bidding documents, and in accordance with the Terms of Reference
Our Proposal shall be valid and remain binding upon us for the period of time specified in the Bid requirements.
I, the undersigned, certify that I am duly authorized by WHO ITB 2023/EURO/MDA/0022 to sign this Proposal and bind it should WHO accept this Proposal.
Name: _____________________________________________________________
Title: _____________________________________________________________
Date: _____________________________________________________________

## LOT 2
Annex 4, LOT 2
ITB 2023/EURO/MDA/0022
Procurement of laboratory consumables and devices, for the National Agency for Public Health
Price Schedule Form
#No | Item description | Technical Specification | Technical specification (supplier proposal) | UM | Quantity | Price/unit USD, VAT exclusive | Amount, USD | Comments
Laboratory Consumables
1 | Calibration plate Applied Biosystems | Calibration plate Applied Biosystems Step One Plus RT PCR System, 96-well optical reaction plates preloaded with dyes (FAM™/SYBR™ Green I, VIC™/JOE™, NED™/TAMRA™/ROX™) are provided for quick and easy calibration of the StepOnePlus™ Real-Time PCR System | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | piece | 1 | 0 | 0 | please provide earliest delivery time, but not later than 31.08.2023
2 | Calibrator set for Quantstudio 5 | Calibration kit establishes the pure dye spectra and multicomponent values for FAM™, VIC™, ROX™, SYBR™ , ABY™, JUN™, MUSTANG PURPLE™, TAMRA™, NED™, and Cy®5 dyes on the QuantStudio™ 5, 96-well 0.2-ml real-time PCR systems. Concentration - 0.2 mL, GC-Rich PCR Performance - Low, PCR Method - qPCR, Volume - 0.2 mL, For Use With (Application) - Real Time PCR (qPCR), Product Type - Spectral Calibration Kit, Shipping Condition - Dry Ice. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | piece | 4 | 0 | 0
3 | ADN/ARN Vibrio cholerae (PCR) RT, calitativ | Reaction type - chain polymerization reaction (PCR) with real-time detection, test for qualitative detection ADN/ARN Vibrio cholerae in the biological material and environmental samples using real-time hybridization-fluorescence detection of amplified products.Reagents for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | pcs | 2 | 0 | 0
4 | Corynebacterium diphtheriae / tox-genes-PCR kit | Reaction type - chain polymerization reaction (PCR) with real-time detection, test for qualitative detection of Corynebacterium diphtheriae DNA and genes encoding toxins of Corynebacterium diphtheriae and Corynebacterium ulcerans in the biological material (nasopharyngeal swabs, oropharyngeal swabs, swabs from disease sites, germ culture), using real-time hybridization-fluorescence detection of amplified products, IVD.Reagents for at least 50 tests, including controls, compatible with QuantStudio 5 amplifier. The set will include reagents for all stages: extraction, amplification with real-time detection. EC certificate for in vitro diagnostic use in accordance with the requirements of the European Directive. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | pcs | 1 | 0 | 0
5 | Cholera Ag test is a rapid immunochromatographic assay for qualitative detection of Vibrio cholera O1/O139 in human fecal | Kit for 20 tests Sensitivity: 100% (O1), 100% (O139) Specificity: 100% (O1), 100% (O139) | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | kit | 12 | 0 | 0
6 | Antitoxin diagnostic diphtheria purified by fermentolysis and specific sorption dry | Kit = 10 ampoules in a cardboard box, the reagent is lyophilized in ampoules with a volume of 1 ml, has the form of an amorphous mass in the form of a white or light yellow pill. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | kit | 12 | 0 | 0
7 | Antimicrobial susceptibility plate for testing Salmonella and E. coli isolates | kit= N10 Testing plate for Salmonella andi E.Coli EUVSEC 3 | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | kit | 2 | 0 | 0
8 | Antimicrobial Susceptibility Plate for Testing Gram-Negative Isolates | kit= N10 Testing plate GN EUMDRXXF | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | kit | 2 | 0 | 0
9 | Antimicrobial Susceptibility Plate for Testing Pseudomonas and Acinetobacter Isolates | kit= N10 Testing plate Pseudomonas andi Acinetobacter EUX2NF | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | kit | 2 | 0 | 0
10 | Mueller-Hinton Broth | Set= N100 Sensititre MH Broth T3462 | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | set | 1 | 0 | 0
11 | Sterile Water | Set= N100 Sensititre Sterile Water T3339 | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | set | 1 | 0 | 0
12 | Dosing heads N100 | set = N100 Dosing heads | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | set | 1 | 0 | 0
13 | Rapid immunochromatographic test C. difficile GDH+Toxin A+B, combo card for the qualitative detection of glutamate dehydrogenase (GDH), toxin A and toxin B in C. difficile in feces. | "Piece=kit/ 20 tests Detection limit from: 0.10 ng/mL C. difficile GDH, 0.25 ng/mL C. difficile Toxin A, 0.19 ng/mL C. difficile Toxin B Positive control GDH, Tox A/B - present in the kit. Sensitivity >97% and specificity >95%, result in maximum 15 min. " | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | kit | 3 | 0 | 0
14 | Clostridium Difficile GDH ELISA Assay Kit | Type of reaction - immunoenzymatic for the qualitative and quantitative determination of GDH C. difficile toxin in feces. The presence in the kit of all reagents required for the reaction, including controls, blank. Plate of 96 wells (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 95.8%, diagnostic specificity not less than 98.9%. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | kit | 5 | 0 | 0
15 | Clostridium Difficile Toxin AB Qualitative ELISA Assay Kit | Type of reaction - immunoenzymatic for the qualitative determination of C. difficile toxins A and B in feces. The presence in the kit of all reagents required for the reaction, including controls, blank. Plate of 96 wells (12 strips of 8 removable wells each) Reagents for 96 investigations including controls. Test sensitivity not less than 98%, diagnostic specificity not less than 91.8%. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | kit | 3 | 0 | 0
16 | Ser anti-V. cholerae /Ogava | Dehydrated, Piece = bottle. Packaging=bottle. min - 1 ml for ready-to-use product, according to order Ministry of Health no. 701 of 18.10.2010 Annex 6 p.9. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | bottle | 40 | 0 | 0
17 | Ser anti-V. cholerae /Inaba | Dehydrated, Piece = bottle. Packaging=bottle. min - 1 ml for ready-to-use product, according to order Ministry of Health no. 701 of 18.10.2010 Annex 6 p.9. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | bottle | 40 | 0 | 0
18 | Ser anti-V. cholerae O139 (Bengal) | Dehydrated, Piece = bottle. Packaging=bottle. min - 1 ml for ready-to-use product, according to order Ministry of Health no. 701 of 18.10.2010 Annex 6 p.9. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | bottle | 40 | 0 | 0
19 | Ser anti-V. cholerae O1 (poly) | Dehydrated, Piece = bottle. Packaging=bottle. min - 1 ml for ready-to-use product, according to order Ministry of Health no. 701 of 18.10.2010 Annex 6 p.9. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | bottle | 40 | 0 | 0
20 | Ser anti-V. cholerae RO | Dehydrated, Piece = bottle. Packaging=bottle. min - 1 ml for ready-to-use product, according to order Ministry of Health no. 701 of 18.10.2010 Annex 6 p.9. | Please do not copy requirements, please provide brand and technical parameters of the equipment proposed, including brochures, ISO certificates and CE Certificates | bottle | 40 | 0 | 0
| TOTAL LOT 2 Laboratry Consumables | [5 empty] | 0
We declare that all the information and statements made in this Proposal are true and we accept that any misinterpretation or misrepresentation contained in this Proposal may lead to our disqualification .
We offer to provide services in conformity with the Bidding documents, and in accordance with the Terms of Reference
Our Proposal shall be valid and remain binding upon us for the period of time specified in the Bid requirements.
I, the undersigned, certify that I am duly authorized by WHO ITB 2023/EURO/MDA/0022 to sign this Proposal and bind it should WHO accept this Proposal.
Name: _____________________________________________________________
Title: _____________________________________________________________
Date: _____________________________________________________________