## Sheet1
ITB 2023/EURO/MDA/0027
Annex 1. List of requirements and specifications
Procurement of the medical Equipment for the Insitute of Oncology
Electrosurgical energy platform unit with accessories (advanced energy and electrosurgery devices)
| Description | Electrosurgical energy platform, providing advanced energies in sealing and cutting vessels, intended for open surgical and laparoscopic procedures, based on energy generation, for urology, gynecology, colorectal, cardiac, ENT and general application.
| Parameters |  |  | Minimum expected Specification | Proposed Specificacion (to be completed by the bidder) | Reference document / brochure / page where provided information may be verified by the evaluation committee
|  | Monopolar with international 3-pin connection |  | 2 outlets
| Output channels | Bipolar and vessel sealing |  | 2 outlets
|  | Hand switch |  | Yes
|  | Bipolar footswitch and vessel sealing |  | Yes
|  | Monopolar footswitch |  | Yes
| Working modes | Monopolar | Pure cut | ≥300W
|  |  | Blend cut | ≥ 200 W
|  |  | Coagulation Lightning | ≥ 120W
|  |  | Coagulation Soft | ≥120 W
|  |  | Hemostasis spray | ≥ 100W
|  | Bipolar | Pure cut (resection or vaporization) | ≥ 370W
|  |  | Coagulation cut (resection or vaporization) | ≥150 W
|  |  | Coagulation | ≥ 90W
|  | Vessel Sealing (Tissue Sensing Technology, Intelligent Tissue Monitoring or equivalent performances) | Technology for reads tissue | Yes
Sealing blood and lymphatic vessels up to and including 7 mm in diameter and tissue bundles, with minimal thermal spread and tissue burns
| Real-time measurement technology of tissue changes | Tissue Sensing Technology control 434000 data points per second (LigaSure); Surgical Tissue Management System – ITM Intelligent Tissue Monitoring (Thunderbeat); or equivalent performances
| Work schedules | Surgical applications |  | minimum 60 programs
| Screen (screens) | LCD touchscreen with the ability to control at least 4 different working modes (Monopolar, Bipolar, Sealing vessels, Laparoscopic) |  | Yes
|  | Separate display in minim 4 quadrants |  | Yes
| Self-test function |  |  | Yes
| Indicators |  |  | acoustic
visual
| Sound volume control |  |  | Yes
| Built-in software for bipolar resection in saline environment |  |  | Yes
| Supply voltage |  |  | 220 - 240 V, 50 Hz
Additional functions
"Plug-and-play" technology, the work sockets are of the "SMART" type, they automatically recognize the work tools and have automatic energy configuration
Automatic recognition of completion of the Vessel Sealing cycle and cycle completion sound
Interactive error communication, including on-screen corrective actions in case of non-completion of the Vessel Sealing cycle
Accessory
| Power cable |  |  | 2 pcs.
The accessories below are produced by the same manufacturer as the energy platform
| Mobile cart with wheels for transporting the Electrosurgical energy platform unit |  |  | 1 pc.
| Single-circuit footswitch for bipolar control, with connection cable included |  |  | 2 pcs.
| 2-circuit footswitch for monopolar/bipolar control with connection cable included |  |  | 2 pcs.
| Reusable neutral electrode, with minimum 3 m cable included |  |  | 5 pcs.
| Reusable monopolar handle (pencil) for the electrode, with at least two operating buttons, min. 3 m cable, minimum 50 autoclavable cycles |  |  | 12 pcs
| Set of electrodes (min 3 pcs.) of different sizes and shapes (ball, loop, blade, needle), autoclavable |  |  | 5 set
| Monopolar HF cable of at least 3 m length (for connecting classical instruments), autoclavable |  |  | 2 pcs.
| Bipolar HF cable of at least 3 m length (for connecting classical instruments), autoclavable |  |  | 2 pcs.
| Bipolar forceps bayonet, non-stick, length 18-26 cm, reusable, autoclavable, with cable at least 3 meters |  |  | 4 pcs.
| Multi-functional instrument for cutting and sealing blood vessels, used in open surgery, length 20-25 cm, diameter 5 mm, curved jaws, cut-seal length 16-20 mm, with interconnecting cable at least 3 meters, disposable |  |  | 30 pcs.
| Multi-functional instrument for cutting and sealing blood vessels, used in open surgery, length 20-25 cm, diameter 5 mm, (straight) blunt jaws, cut-seal length 16-20 mm, with interconnecting cable at least 3 meters, disposable |  |  | 20 pcs.
| Multi-functional instrument for cutting and sealing blood vessels, used in open surgery, length 20-25 cm, diameter 9-10 mm, large jaws, with interconnecting cable at least 3 meters, disposable |  |  | 20 pcs.
| Multi-functional instrument for cutting and sealing blood vessels, used in laparoscopy, length 35-40 cm, diameter 5 mm, wide curved jaws, cut-seal length 16-20 mm, with interconnecting cable at least 3 meters, disposable |  |  | 6 pcs.
Other requirements
| 13485 standard, CE certification, Standards and classifications: meets all the conditions required by the standards IEC 60601-1 edition 2.0 and 3.1, IEC 60601-2-2 edition 4.0 and 5.0, IEC 60601-1-2 edition 2.1 and 3.0 and 4.0, IEC 60601-1-8 edition 2.1. |  |  | required
| The platform offers a high degree of protection against electric shocks and can be used in cardiac procedures CF Equipment/Defibrillator Proof (IEC 60601-1, IEC 60601-2-2, and ANSI/AAMI HF18) |  |  | required
| Electrical safety Class 1, CF type |  |  | required
| Year of production minimum 2022 |  |  | required
| Preventive maintenance services during the warranty period. |  |  | required
| In case of failure of the device under warranty, the repair time will be added to the full warranty period. During the warranty period, the intervention time in case of failure of at least 36 hours. |  |  | required
| User manual, service manual, according to law 102 of 2017 with all amendments., Installation, commissioning mandatory |  |  | required
| Installation, commissioning, training by the winning participant - mandatory |  |  | required
| Training for users on installation and on request - mandatory |  |  | required
| User manual in Romanian and English |  |  | required
| Minimum 24 months warranty from the moment of commissioning. |  |  | required