## A3 Telemedicine cart
Annex 1
ITB 2024/EURO/MDA/0001
Procurement of the Telemedicien Cart Units for the STROKE centers within the country hospitals
List of requirements and specifications
NAME, CATEGORY AND CODING
1 | Quantity | 13
2 | Generic name | Telemedicine cart
3 | Level of use | District hospital, raion hospital, national hospital, specialized hospital.
4 | Place of use | Intensive Care Unit, Emergency, Recovery, Medical Imaging, Outpatient Services, Inpatient Care, Specialized Diagnostics, Specialized Treatment
PURPOSE OF USE
1 | Clinical or other purpose | Medical consultation services through 2 key hospitals, which will receive consultations to diagnose patients from 11 primary stroke centres. This system requires: -11 telemedicine cart mobile stations for primary stroke centres -2 fixed telemedicine systems to receive calls from the mobile telemedicine stations.
2 | Overview of functional requirements | This will consist of mobile telemedicine stations that will be able to communicate with the fixed telemedicine stations where the specialist will receive consultations. The mobile telemedicine stations will contain all the medical equipment necessary for the medical consultation. All telemedicine stations should be able to connect to each other and exchange information.
Yes
No
Item | Group | Category | Description | Acceptability Criteria | Meet WHO SPECS | Proposed Specificacion (to be completed by the bidder) | Reference document / brochure / page | Unknown
1 | General Technical Characteristics | Global System Composition | Fixed telemedicine station - 2 pcs | Mandatory
2 | General Technical Characteristics | Global System Composition | Software - for the entire system | Mandatory
3 | General Technical Characteristics | Global System Composition | Mobile telemedicine station with interconnected medical - 11 pcs | Mandatory
4 | General Technical Characteristics | Global System Composition | The following specifications specify the composition of each Global System Composition | Mandatory
5 | Fixed telemedicine station at hospital | Base Cart (if available) | Height Adjustment 70-100 cm | Mandatory
6 | Fixed telemedicine station at hospital | Base Cart (if available) | Display holder, Vesa compatible Monitor Mount | Mandatory
7 | Fixed telemedicine station at hospital | Base Cart (if available) | Casters: medical casters 125mm diameter approx., min 2 locking | Mandatory
8 | Fixed telemedicine station at hospital | Base Cart (if available) | Locking Supply Drawers [2] | Important
9 | Fixed telemedicine station at hospital | Base Cart or area work | Footprint 19.7" x 18.5" (500 x 470 mm) approx. | Important
10 | Fixed telemedicine station at hospital | Base Cart or area work | Mounts for video or web camera above the monitor | Mandatory
11 | Fixed telemedicine station at hospital | Base Cart or area work | Work Surface min 18.5" x 19.3" | Important
12 | Fixed telemedicine station at hospital | Display | Min 24" wide TFT color LCD | Mandatory
13 | Fixed telemedicine station at hospital | Display | Resolution: Min. 1920 x 1080 FHD | Mandatory
14 | Fixed telemedicine station at hospital | Computing System | CPU - min. 12 Cores, Base Frequency min. 3.5 GHz | Mandatory
15 | Fixed telemedicine station at hospital | Computing System | RAM - min. 16GB, DDR4 | Mandatory
16 | Fixed telemedicine station at hospital | Computing System | HDD/SSD Storage - at least 1TB | Mandatory
17 | Fixed telemedicine station at hospital | Computing System | Computer Keyboard & Mouse [1] | Mandatory
18 | Fixed telemedicine station at hospital | Computing System | WLAN and bluetooth | Mandatory
19 | Fixed telemedicine station at hospital | Computing System | Supports Microsoft Windows10; Windows 11 | Mandatory
20 | Fixed telemedicine station at hospital | Computing System | Peripheral streaming - mobile camera | Mandatory
21 | Fixed telemedicine station at hospital | Computing System | HDMI port integrated | Mandatory
22 | Fixed telemedicine station at hospital | Computing System | USB 2.0 ports [≥3] | Mandatory
23 | Fixed telemedicine station at hospital | Computing System | Authentication by fingerprint or password for access | Mandatory
24 | Fixed telemedicine station at hospital | Audio | Built in loudspeakers | Mandatory
25 | Fixed telemedicine station at hospital | Audio | Echo cancellation | Mandatory
26 | Software | Detailed requirements | Work offline when there's no internet connection, and upload the information once internet access is restored. | Mandatory
27 | Software | Detailed requirements | Software for remote medical consultations | Mandatory
28 | Software | Detailed requirements | Integrated medical agenda | Mandatory
29 | Software | Detailed requirements | Patient information display and modification | Mandatory
30 | Software | Detailed requirements | Data collection of patient health parameters | Mandatory
31 | Software | Detailed requirements | Report analysis | Mandatory
32 | Software | Detailed requirements | Remote configuration | Mandatory
33 | Software | Detailed requirements | DICOM compatible | Mandatory
34 | Software | Detailed requirements | HL7 or EMR compatible | Important
35 | Software | Detailed requirements | Ability to interface with hospital patient record software | Mandatory
36 | Software | Detailed requirements | One-time purchase software (or specify) | Important
37 | Software | Detailed requirements | Software for 15 users | Mandatory
38 | Software | Detailed requirements | Romanian lenguage | Important
39 | Mobile telemedicine station | System Composition | Mobile telemedicine station with interconnected medical devices: 11 units | Mandatory
40 | Mobile telemedicine station | System Composition | Cart with castors or station | Mandatory
41 | Mobile telemedicine station | System Composition | Computer w/ keyboard & mouse | Mandatory
42 | Mobile telemedicine station | System Composition | General Examination Camera | Mandatory
43 | Mobile telemedicine station | System Composition | ECG 12 channels | Mandatory
44 | Mobile telemedicine station | System Composition | Blood Pressure monitor | Mandatory
45 | Mobile telemedicine station | System Composition | Oximeter- SpO2 | Mandatory
46 | Mobile telemedicine station | System Composition | Digital Thermometer | Mandatory
47 | Mobile telemedicine station | System Composition | Ultrasound Doppler | Mandatory
48 | Mobile telemedicine station | System Composition | Digital Stethoscope | Mandatory
49 | Mobile telemedicine station | System Composition | UPS | Mandatory
50 | Mobile telemedicine station | Cart with castors or station | Height Adjustment | Mandatory
51 | Mobile telemedicine station | Cart with castors or station | Display holder, Vesa compatible Monitor Mount | Mandatory
52 | Mobile telemedicine station | Cart with castors (if available) | Casters: medical casters 125mm diameter approx., min 2 locking | Mandatory
53 | Mobile telemedicine station | Cart with castors or station | Locking Supply Drawers [2] | Important
54 | Mobile telemedicine station | Cart with castors or station | Dedicated space in telemedicine cart for the medical equipment | Mandatory
55 | Mobile telemedicine station | Cart with castors or station | Footprint 19.7" x 18.5" (500 x 470 mm) approx. | Important
56 | Mobile telemedicine station | Cart with castors or station | Mounts for video or web camera above the monitor | Mandatory
57 | Mobile telemedicine station | Cart with castors or station | Work Surface min 18.5" x 19.3" | Important
58 | Mobile telemedicine station | General Examination Camera | Resolution - min. full HD 1080p 60fps | Mandatory
59 | Mobile telemedicine station | General Examination Camera | Total zoom - at least 18X | Important
60 | Mobile telemedicine station | General Examination Camera | Pan >120° | Mandatory
61 | Mobile telemedicine station | General Examination Camera | Tilt >20° | Mandatory
62 | Computer w/ keyboard & mouse | Display | >15" wide TFT color LCD | Mandatory
63 | Computer w/ keyboard & mouse | Display | Resolution: Min. 1920 x 1080 FHD | Mandatory
64 | Computer w/ keyboard & mouse | Computing System | Min. 12 Cores, Base Frequency min. 3.5 GHz | Mandatory
65 | Computer w/ keyboard & mouse | Computing System | RAM - min. 16GB, DDR4 | Mandatory
66 | Computer w/ keyboard & mouse | Computing System | HDD/SSD Storage - at least 1TB | Mandatory
67 | Computer w/ keyboard & mouse | Computing System | WLAN and bluetooth | Mandatory
68 | Computer w/ keyboard & mouse | Computing System | Supports Microsoft Windows10; Windows 11 | Mandatory
69 | Computer w/ keyboard & mouse | Computing System | Real time Screen sharing and live annotation in remote area and with the specialist | Mandatory
70 | Computer w/ keyboard & mouse | Computing System | Audio and video recording capable | Mandatory
71 | Computer w/ keyboard & mouse | Computing System | Built in loudspeakers for teleconsult | Mandatory
72 | Computer w/ keyboard & mouse | Computing System | Echo cancellation | Mandatory
73 | ECG 12 channels | Detailed requirements | 12 channel ECG with 12 leads | Mandatory
74 | ECG 12 channels | Detailed requirements | Limb derivations: I, II, III, aVL, aVF, aVR | Mandatory
75 | ECG 12 channels | Detailed requirements | Precordial derivations: V1, V2, V3, V4, V5, V6 | Mandatory
76 | ECG 12 channels | Detailed requirements | Manual selection for any 12-lead | Important
77 | ECG 12 channels | Detailed requirements | Respiratory Rate: 0 - 150 rpm | Mandatory
78 | ECG 12 channels | Alarms | Lead Disconnect | Mandatory
79 | ECG 12 channels | Alarms | Low battery | Mandatory
80 | ECG 12 channels | Electrical requirements | Operating time >2 hours | Mandatory
81 | ECG 12 channels | Electrical requirements | Charging time <7 hours | Mandatory
82 | ECG 12 channels | Software | ST, QT & arrhythmia analysis | Mandatory
83 | ECG 12 channels | Software | Pacemaker impulse detection | Mandatory
84 | ECG 12 channels | Accessories | ECG cable for 12 lead [1] | Mandatory
85 | ECG 12 channels | Accessories | Reusable set of Chest Suction and Limbs Electrodes set [2] | Mandatory
86 | ECG 12 channels | Accessories | Conductive gel 100 ml [15] | Mandatory
87 | Blood Pressure monitor | Accessories | Adult StO2 cuff [2] | Mandatory
88 | Blood Pressure monitor | Accessories | Pediatric StO2 cuff [2] | Mandatory
89 | Oximeter- SpO2 | Detailed requirements | Saturation range: 0-100% | Mandatory
90 | Oximeter- SpO2 | Detailed requirements | Pulse range 20-250 bpm | Mandatory
91 | Oximeter- SpO2 | Detailed requirements | SPO2 accuracy: + / - 2% | Mandatory
92 | Oximeter- SpO2 | Detailed requirements | Pulse rate accuracy: + / - 1 BP | Mandatory
93 | Oximeter- SpO2 | Electrical requirements | Battery with internal charger unit | Mandatory
94 | Oximeter- SpO2 | Electrical requirements | Battery operating time > 4 hours | Mandatory
95 | Digital Thermometer | Detailed requirements | Led display | Mandatory
96 | Digital Thermometer | Detailed requirements | Gun Shape, or similar | Mandatory
97 | Ultrasound Doppler | Transducers | Transducer continuous wave Doppler (detect and characterize blood flow) | Mandatory
98 | Ultrasound Doppler | Modes | Color Doppler Imaging | Mandatory
99 | Ultrasound Doppler | Modes | Pulsed wave Doppler | Mandatory
100 | Ultrasound Doppler | Software | Vascular | Mandatory
101 | Ultrasound Doppler | Software | Blood flow abnormalities | Mandatory
102 | Ultrasound Doppler | Accessories | Conductive gel 100 ml [15] | Mandatory
103 | Digital Stethoscope | General Technical Characteristics | Wireless | Mandatory
104 | Digital Stethoscope | Alarms | Low battery | Mandatory
105 | General Technical Characteristics | Electrical requirements | Power 220 V, 50 Hz. Electrical power connector should meet German standard Shuko type | Mandatory
106 | General Technical Characteristics | Electrical requirements | Battery time of the entire system > 4 hrs | Mandatory
107 | General Technical Characteristics | Electrical requirements | Charging time <6 hrs | Mandatory
108 | General Technical Characteristics | Electrical requirements | All the complementary equipment must include the charger | Mandatory
109 | Warranty and maintenance | Service Conditions | Period: at least 24 months | Mandatory
110 | Warranty and maintenance | Service Conditions | Maximum response time: 72 hours | Mandatory
111 | Warranty and maintenance | Service Conditions | Preventive maintenance tasks included in the warranty twice a year | Mandatory
112 | Warranty and maintenance | Service Conditions | Software upgrade in warranty | Mandatory
113 | Installation |  | Supplier must perform the installation, safety and operation checks before handover | Mandatory
114 | Installation |  | Equipment assembly and commissioning on site at beneficiary location | Mandatory
115 | Transportation |  | Supplier must include the transportation to the final destination | Mandatory
116 | Training and utilisation |  | Training for all medical personnel who will be using this item, on each individual site | Mandatory
117 | Training and utilisation |  | Training for at least 2 technical personnel on each site at Beneficiary location | Mandatory
118 | Training and utilisation |  | Training provided on each individual site | Mandatory
119 | Training and utilisation |  | Training in Romanian or Russian language (if not compliant specify) | Mandatory
120 | Training and utilisation |  | Training on basic features and functions | Mandatory
121 | Training and utilisation |  | Training on advanced features and functions | Mandatory
122 | Training and utilisation |  | Training on daily maintenance and cleaning procedures | Mandatory
123 | Training and utilisation |  | Training on common faults/problems detection and solving | Mandatory
124 | Training and utilisation |  | Training on preventive maintenance every 6 months | Mandatory
125 | Training and utilisation |  | Duration of the training (specify) | Mandatory
126 | Environmental requirements |  | Capable of being stored continuously in ambient temperature range at least from 0 to 50 deg C and relative humidity from 15 to 90%. Capable of operating continuously in ambient temperature range at least from 10 to 40 deg C and relative humidity from 15 to 90%. | Mandatory
127 | Documentation |  | All supporting documentation must be presented at least in English and, when available, in Romanian or Russian | Mandatory
128 | Documentation |  | Datasheet or manual of each medical device (general examination camera, ECG 12 channels, blood pressure monitor, oximeter- SpO2, digital thermometer, US doppler, digital stethoscope, UPS ) | Mandatory
129 | Documentation |  | Manual procedures required for local calibration and routine maintenance | Mandatory
130 | Documentation |  | Troubleshooting protocols | Mandatory
131 | Documentation |  | Cleaning and disinfection protocols | Mandatory
132 | Documentation |  | List of spares and accessories, with their part numbers and cost. | Mandatory
133 | Documentation |  | Two user’s manuals, preferable in Romanian language (also in electronic format) | Mandatory
134 | Documentation |  | Two service manual, preferable in Romanian language (also in electronic format) | Mandatory
135 | Local representative/partner | To provide service for the product | Located in Moldova/Romania (specify address) | Mandatory
136 | Local representative/partner | To provide service for the product | Name | Mandatory
137 | Local representative/partner | To provide service for the product | Street / N° | Mandatory
138 | Local representative/partner | To provide service for the product | ZIP / City | Mandatory
139 | Local representative/partner | To provide service for the product | Country | Mandatory
140 | Local representative/partner | To provide service for the product | Contact person and phone number | Mandatory
141 | Local representative/partner | To provide service for the product | Years of experience in medical equipment procurement | Mandatory
142 | Local representative/partner | To provide service for the product | Number of technicians skilled for the offered equipment (specify) | Mandatory
143 | Safety and Standards | European standards | (EU) 2016/679 General Data Protection Regulations (GDPR) | Mandatory
144 | Safety and Standards | European standards | 2017/745/EEC for each MD class II | Mandatory
145 | Safety and Standards | International standards | ISO 13485:2016 Medical devices — Quality management systems — Requirements for regulatory purposes | Mandatory
146 | Safety and Standards | International standards | ISO 14001:2015 Environmental management systems — Requirements with guidance for use | Mandatory
147 | Safety and Standards | International standards | IEC 60601-1 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. | Important
148 | Safety and Standards | International standards | ISO/TS 13131:2014 Health informatics — Telehealth services — Quality planning guidelines | Important
149 | Safety and Standards | Regulations | Free Sales Certificate (FSC) provided by any of the following countries: Australia, Canada, Japan, USA and/or European Community or Declaration of Conformity | Important
150 | Safety and Standards | Regulations | All signs (on monitor displays, switches, buttons, labels, keyboard) should be in Romanian or English | Mandatory
151 | Safety and Standards | Regulations | Manufacturer`s authorization | Mandatory

## 16-Dispo.padeles ad.ped. (hidden)
Item Disposable padeles for adults and pediatricc