## WHO_MDA_ Autoclave techspecs
Annex
2024/EURO/MDA/0012a
Steam sterilizer
NAME, CATEGORY AND CODING
Parameters |  |  | Minimum expected Specification | Proposed Specificacion (to be completed by the bidder) | Reference fom manufacturer original document / brochure / page where provided information may be verified by the evaluation committee
1 | Generic name | Steam Sterilizer 90l, class B
PURPOSE OF USE
2 | Clinical or other purpose | The unit is designed to cover a large field of applications for hospital CSSDs, pharmaceutical and biotechnological applications. Fully automatic high-speed pre & post vacuum steam sterilizer. The model is an electrically heated sterilizer, which operates with saturated steam as a sterilizing agent, and has a temperature between 105°C and 138ºC and pressure up to 3 bars. The autoclave is classified as Class B according to MDD 93/42/EEC. The autoclave is designed as a Large Steam Sterilizer in accordance with EN285. The device is not intended for use in the presence of flammable mixtures. | required
3 | Level of use | District hospital, provincial hospital, national hospital, specialized hospital. | yes
4 | Clinical department/ward | Decontamination, Sterilization | yes
TECHNICAL CHARACTERISTICS and PHYSICAL CHARACTERISTICS
5 | CAPACITY | Chamber volume (conventional 90l), horisontal, front-loading, | min. 85l-100l max.
6 | FRAME, BODY | Hinged door with automatic locking or Manual door, will be accepted | required
|  | AISI 316L stainless steel chamber | required
|  | The design can be as standalone or tabletop(descktop) version | required
|  | Combination of painted quality steel cover and plastic for front panel or AISI 304L stainless steel front, side cover | required
|  | piping, pumps and fittings are made from 316L stainless steel | required
|  | The sealing of the door must be ensured by silicone gasket | required
|  | The door must be designed with a safety system to prevent injury to the user | required
|  | Type of sterilization chamber- rectangular or round | required
7 | STEAM GENERATOR | Built-in electrical steam generator. The sterilizer has a fully-automatic built-in steam generator, with electric elements as the heat source. | required
|  | Internal heaters with separate control and short-circuit safety | required
|  | Overheat protection. This protection prevents the steam generator from overheating if the boil-dry protection should fail to operate. | required
|  | Pressure equalization in the sterilization chamber ensured by HEPA filter | required
|  | AISI 316L stainless steel generator chamber | required
8 | CICLE | Preset sterilization regimes | >5 programms
|  | Sterilization regimes that can be programmed by the user | >5 programms
|  | Setting by the user of the desired sterilization regime | required
|  | Tests supported: vacuum test (leakage) and Bowie-Dick test | required
|  | Electronic control: microprocessor | required
|  | Digital monitoring: chamber temperature, chamber pressure, time, sterilisation status | required
|  | Equipped with water-based vacuum pump, integrated | required
|  | Display of the sterilization graph on the screen | required
|  | Operation up to 140ºC - 2.5bar | required
|  | End of stery cicle must have a printing report on paper (parallel storing data in memory) | required
9 | SAFETY | Over temperature protection | required
|  | Alarms: power failure, low temperature, sterilization cycle failure | required
|  | The chamber will maintain a test pressure ≥3 bar, equipped with safety valve | required
|  | Doors must be thermally insulated to prevent the surface temperature presenting a potential hazard to operators | required
|  | filtration systems ensure air is decontaminated before discharge | required
|  | Safety door opening system in case of pressure in the sterilization chamber | required
ELECTRICAL CHARACTERISTICS
10 |  | 220V, 50Hz | required
Electronical
11 | Electronical | minim 5" multi-color screen panel | required
|  | Internal memory for storing minim 500 cycles | required
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
12 |  | Grease for gasket minim 100gr | required
|  | maintenance kit for first year(must include but not only, a door gasket, HEPA filter) | required
|  | integrated Data Printer (paper used on printer must have universal dimensions used in other units) | required
|  | loading rack, in 2 levels variant or shelves in 2 levels | required
|  | Printer paper 10 pcs | 10 pcs
|  | Air Compressor (oil free), in case if the sterilizer use compressed aIr | optional
TRAINING, INSTALLATION AND UTILISATION
13 | Transportation | Supplier must include the transportation to the final healthcare facility | required
14 | Installation | Supplier must perform the fully installation, safety and operation checks before handover. All operations must have a conformities reports. Training for users and techicians must be provided. | required
WARRANTY AND MAINTENANCE
15 | Warranty | minim 24 months | required
DOCUMENTATION
16 | Documentation requirements | All supporting documentation, Technical/Maintenance manuals to be supplied in English. User manual in romanian language, according with LAW No.102 from 09-06-2017 regarding medical devices including: | required
SAFETY AND STANDARDS
17 | Regulatory Approval / Certification | Free Sales Certificate (FSC) provided by any of the following countries: Australia, Canada, Japan, USA and/or European Community (e.g. CE certificate given by a third certified party for the specific medical devices proposed (not only declaration of conformity)). | required
18 | Standards for the manufacturer | all valid certificates listed below: • CE marking certificate the EU (EC ) declaration of conformity) • Certified quality management system for medical devices- ISO 13485 • ISO 9001:2015 - Quality Management Systems • IEC 61010- Safety requirements for sterilizers used to treat medical materials | all sertificates must be presented in copies with confirmation stamp

## WHO_MDA_ pumps techspecs
Annex 1
Syringe Pump (Single Channel)
NAME, CATEGORY AND CODING
Parameters |  |  | Minimum expected Specification | Proposed Specificacion (to be completed by the bidder) | Reference fom manufacturer original document / brochure / page where provided information may be verified by the evaluation committee
1
2 | Generic name | Syringe Pump
3 | Specific type or variation
4 | GMDN name | Syringe pump
5 | GMDN code | 13217
6 | GMDN definition | A mains electricity (AC-powered) device designed to precisely drive the plunger of a syringe down its barrel to infuse a solution when it must be administered with a high degree of volume accuracy and rate consistency. Because of the lower flow settings and flow resolution (e.g., 0.1 ml/hr), it is especially appropriate for neonatal, infant, and critical care applications in which small volumes of concentrated drugs are to be delivered over an extended period. It can also be used to administer epidural analgesia. It will typically have internal batteries that allow the device to operate for a short period of time when no line power is available (e.g., during transport or a power outage).
PURPOSE OF USE
7 | Clinical or other purpose | Syringe pump is a device that allows the intravenous administration of medical fluids with high accuracy.
| Level of use | District hospital, provincial hospital, national hospital, specialized hospital.
| Clinical department/ward | Intensive Care Unit, Emergency, Recovery, Medical Imaging, Outpatient Services, Inpatient Care, Specialized Diagnostics, Specialized Treatment
| Overwiew of functional requirements | A syringe pump, is as a syringe driver, that helps deliver an accurate amount of fluids, whether nutrients or medications, into the patient’s body in a controlled manner. They are used in medical settings such as a clinic, nursing home, or hospital, but also during research studies. A syringe pump is a motor-driven device that has inbuilt software to program the rate and duration of the fluid. A syringe infusion pump has a step motor that pushes the syringe plunger that ejects the fluids through the needle into the patient’s body.
8 | GENERAL REQUIREMENTS | Designed for long use and frequent and easy dismount anddisinfection with hospital-grade products | required
|  | Fot high accuracy and small flow rate liquid transferring | required
|  | Syringe Pump with open system option: accepts all standard syringes of 2, 3, 5, 10, 20, 30, 50, 60ml. | required
|  | MINIM Display settings-display volume, flow rate or linear speed, speed, date, time, battery status, alarms | required
|  | Memory Function | required
|  | Calibration Function | required
|  | Syringe Protection | required
|  | Ability to add new types of syringes to the library | required
|  | Automatic syringe detection | required
TECHNICAL CHARACTERISTICS
9 | Technical specifications | Work modes, minimum | Rate /Time/Dose/Sequential/Drug library Mode
|  | Syringe Channels | 1
|  | Display LCD/TFT | required
|  | the possibility of adjusting the brightness | required
|  | Flow rate | 0.1 - 999 ml/h
|  | Flow rate step | 0.01 mL/h
|  | Anti-bolus system | required
|  | Bolus Rate | 0.1-2000ml/h
|  | Automatic/Manual bolus available | required
|  | flow rate accuracy | ±2% and better
|  | occlusion pressure level | 75 - 880 mmHg and better
|  | Syringe size settings | minimum 6 pre-sets, alow set customisable diameter
|  | ALARMS displayed , minimal | Syringe Holder Loose, Syringe Flange Insertion Error, Syringe Not Calibrated, Occlusion, Syringe Empty, Rate Abnormal, No Battery, Infusion Complete, Syringe Near Empty, Battery Low
|  | Sound alarms settings | Adjustable volume, Silent mode button
|  | Protection Class | Class I, IP34
|  | Mounting bracket for fixation to standard bed/wallrail and mobile pole/stand. | required
|  | History records | miinim 1000 events
UTILITY REQUIREMENTS
10 | Electrical | The voltage of the electric current | 220, ±10%; 50 Hz
|  | Plug type in compliance with required local/national standards and regulations | required
|  | Appropriate external device to protect the equipment against over-voltage and over-current line conditions (between plug and socket) | required
|  | Rechargeable battery with back up time of minimum 3 hours autonomy | required
TRAINING, INSTALLATION AND UTILISATION
11 | Transportation | Supplier must include the transportation to the final healthcare facility | required
| Installation | Supplier must perform the installation, safety and operation checks before handover | required
| Training of user/s | Training of users in hardware and software (including applications): | required
|  | • Basic operation of the system | required
|  | • Advanced operation of the system | required
|  | • Maintenance routine | required
WARRANTY AND MAINTENANCE
12 | Warranty | minim 24 months | required
| Preventive maintenance | Included in the warranty according with the interval and period recommended by the manufacturer. All nececary consumables and materials needed for maintenance during the given period to be included in the offer. | required
| Software / Hardware upgrade availability | (any avaiable updated during the warranty period) Included | required
DOCUMENTATION
13 | Documentation requirements | All supporting documentation, Technical/Maintenance manuals to be supplied in English. User manual in romanian language, according with LAW No.102 from 09-06-2017 regarding medical devices including: | required
|  | • Manual procedures required for local calibration and routine maintenance | required
|  | • Troubleshooting protocols | required
|  | Full set of requirement documents for registering unit on agency of National Agency for the Regulation of Nuclear and Radiological Activities | required
|  | • Cleaning and disinfection protocols | required
|  | List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Certificate of calibration and inspection to be provided. Contact details of manufacturer, supplier and local service agent to be provided | required
SAFETY AND STANDARDS
14 | Regulatory Approval / Certification | CE certificate/Ec declaration of conformity, Classified under EU MDD 93/42/ECC, or under EUMDR 2017/745 as Class IIb device | required
| Standards for the manufacturer | ISO 13485:2003 Medical devices -- Quality management systems IEC 60601-1:2005 + A1:2012 Medical electricalequipment - Part 1: General requirements for basicsafety and essential performance ISO 9001 Quality management systems | required

## stretcher spec
PATIENT TRANSFER STRETCHER
NAME, CATEGORY AND CODING
Parameters |  |  | Minimum expected Specification | Proposed Specificacion (to be completed by the bidder) | Reference fom manufacturer original document / brochure / page where provided information may be verified by the evaluation committee
1
2 | Generic name | Stretchers
2 | GMDN name | Stretchers, Transfer
3 | GMDN code | 13820
4 | general definition | Transfers are defined as moving a patient from one flat surface to another, such as from a bed to a stretcher. Types of hospital transfers include bed to stretcher, bed to wheelchair, wheelchair to chair, and wheelchair to toilet, and vice versa.
PURPOSE OF USE
5 | Clinical or other purpose | stretcher for transporting seriously ill patients who require permanent monitoring. | required
| Level of use | District hospital, provincial hospital, national hospital, specialized hospital.
| Clinical department/ward | Intensive Care Unit, Emergency, Recovery, Medical Imaging, Outpatient Services, Inpatient Care, Specialized Diagnostics, Specialized Treatment
TECHNICAL CHARACTERISTICS
6 | TYPE AND GENERAL DESCRIPTION | HYDRAULIC PATIENT TRANSFER STRETCHER, 2columns design construction, accessible for C-arm and x-ray | required
|  | stretcher must have minim 2 sections surface | required
|  | Trendelenburg and Reverse Trendelenburg | required
|  | X-Ray transparency | required
|  | X-Ray Cassette holder sliding under mattress platform, minim 2 places | required
|  | Oxygen cylinder holder (5 liters) | required
|  | Retractable and Foldable Stainless Steel side rails with locking mechanism | required
|  | Shockproof protection on four sides | required
|  | 5th Wheel for easy maneuverability | required
|  | Monitor Shelf, with universal fixing system | required
|  | Foam mattress with removable, anti-static and fire-retardant, waterproof and compatible with the size of the stretcher, disinfectant resistibile, 10 cm mattress thickness | required
|  | Backrest adjustment assisted by gas spring | required
|  | Foot operation pedal, available on both sides of stratcher | required
|  | Pushing handles from both sides, foldable | required
7 | General configuration
|  | approximate dimensions, Length,width | 215x82 cm
|  | guaranteed loading capacity | 250kg
|  | Height adjustment, low-hi, minim configuration and better | 65-90 cm
X ray transparent technology platform
|  | special place (holder) for 5 Liter oxygen cylinder wit reduction and flowmeter, attached humidifier, to be installed on stratcher. | required, for 1 cyllinder
|  | Center-locking brake system | required
|  | Dual, side-mounted foot control hydraulics with uni-lower pedal | required
|  | Dual end brake/steer controls | required
|  | Glideaway Siderails | required
|  | Trendelenburg ±16° (±3°) | required
|  | minim 180mm casters with integrated wheel covers | required
|  | IV pole options: Permanent two- or three-stage, Removable | 1 pcs
|  | anti colision wheels all corners | required
|  | Equiped with fifth guide weel | required
|  | Back section adjustment | minimum 0-80°
|  | Mattress, fully x-ray transparent | required
|  | Mattress surface, minim size | 62x194cm
|  | Mattress thickness, | 8-15 cm
|  | Patient Monitor Shelf, with universal fixing system | required, for 1 monitor
TRAINING, INSTALLATION AND UTILISATION
8 | Transportation | Supplier must include the transportation to the final healthcare facility | required
| Installation | Supplier must perform the installation, safety and operation checks before handover | required
| Training of user/s | Training of users | required
|  | • Basic operation of the stratcher | required
|  | • Maintenance routine | required
WARRANTY AND MAINTENANCE
9 | Warranty | minim 24 months | required
| Preventive maintenance | Included in the warranty according with the interval and period recommended by the manufacturer. All nececary consumables and materials needed for maintenance during the given period to be included in the offer. | required
DOCUMENTATION
10 | Documentation requirements | All supporting documentation, Technical/Maintenance manuals to be supplied in English. User manual in romanian language, according with LAW No.102 from 09-06-2017 regarding medical devices including: | required
|  | • Troubleshooting protocols | required
|  | • Cleaning and disinfection protocols | required
SAFETY AND STANDARDS
11 | Regulatory Approval / Certification | Ce certificate or EC declaration of conformity | required
| Standards for the manufacturer | ISO 13485:2003 Medical devices -- Quality management systems ISO 9001:2015 | required

## set oxygen therapy specs
Annex 1
MEDICAL Oxygen Cylinder with REGULATOR and humidifier
NAME, CATEGORY AND CODING
Parameters |  |  | Minimum expected Specification | Proposed Specificacion (to be completed by the bidder) | Reference fom manufacturer original document / brochure / page where provided information may be verified by the evaluation committee
PURPOSE OF USE
1 | Clinical or other purpose | This installation is intended to be installed on the stretcher, for the transport of patients who require assisted ventilation, to be transported inside the medical institution, between different departments. | required
2 | Level of use | District hospital, provincial hospital, national hospital, specialized hospital.
3 | Clinical department/ward | Intensive Care Unit, Emergency, Recovery, Medical Imaging, Outpatient Services, Inpatient Care
4 | Overwiew of functional requirements | The breathing kit will be used to assist patients in transporting them between different departments within the hospital. Medical oxygen cylinder with Oxygen Regulator and Flow Meter and Humidifier Bottle. It has fine adjustment valve with Flowmeter and Humidifier. The metal construction is in superior quality Brass It is suitable for Bull-Nose type valves and Includes flowmeter. The humidifier bottle is made of Polycarbonate which suitable for Oxygen Cylinders
5 | Configuration
|  | Valve type Bull Nose | DIN 477 (G 3/4 ”)
|  | Filled gas | medical oxygen
|  | Type of cylinder | steel (seamless steel)
|  | Bottle paint | white
|  | Working pressure | 200 bar
|  | Test pressure | 300 bar
|  | Capacity (water) | 5l
|  | Service Free over product life time | required
TECHNICAL CHARACTERISTICS
6 | medical cylinder regulator
|  | Adjustable output flow control | 1-25 L/MIN
|  | Pressure gauge | required
|  | Range | 0-300BAR
|  | precisiion | CLASS 2.5
|  | Cylinder connection | DIN 477 (G 3/4 ”)
|  | Fitting | GERMAN (DIN)
|  | Equiped with humidifier and flowmeter | required
|  | Delivery flow (output) | MAX 4BAR
|  | Output connection (Therapy Outlet Connector) | 6MM DIAMETER HOSE CONNECTOR
|  | autoclavable humidifier | up 134 C
|  | Humidifier capacity | min 200 ml + 100ml
|  | Service Free over product life time | required
TRAINING, INSTALLATION AND UTILISATION
8 | Transportation | Supplier must include the transportation to the final healthcare facility | required
9 | Installation | Supplier must perform the installation, safety and operation checks before handover | required
10 | Training of user/s | Training of users | required
WARRANTY AND MAINTENANCE
11 | Warranty | minim 24 months | required
DOCUMENTATION
12 | Documentation requirements | All supporting documentation, Technical manuals to be supplied in English. User manual in romanian language, according with LAW No.102 from 09-06-2017 regarding medical devices including: | required
|  | Cylinder and flow meter, gauge, must have Certificate of calibration and inspection | required
|  | • Cleaning and disinfection protocols | required
required
SAFETY AND STANDARDS
13 | Standards for the manufacturer | ISO 13485:2016 Medical devices -- Quality management system for the design of medical systems ISO 98091:2019 Gas cylinders ISO 9001:2015 ISO 10524-1:2019/A1 Pressure regulators for use with medical gases – Part 1: Pressure regulators and pressure regulators with flow-metering devices | required

## patient mobile monitor
Annex 1
MULTI-PARAMETER PATIENT MONITOR, portable
NAME, CATEGORY AND CODING
Parameters |  |  | Minimum expected Specification | Proposed Specificacion (to be completed by the bidder) | Reference fom manufacturer original document / brochure / page where provided information may be verified by the evaluation committee
1
2 | Generic name | MULTI-PARAMETER Monitor,patient,portable with accessories
3 | GMDN name | Transportable physiologic monitoring system
4 | GMDN code | 36872
5 | GMDN definition | An assembly of battery-powered devices designed for continuous assessment of several vital physiologic parameters (e.g., ECG, blood pressure, heart rate, temperature, cardiac output, apnoea, and respiratory/anaesthetic gas concentrations) of a patient in transit [e.g., during transport to, between, or within a hospital(s)]. It typically includes a recorder, amplifier, display monitor, and patient leads/attachments; it may also be used by the bedside where it is normally connected to the mains electricity.
PURPOSE OF USE
6 | Clinical or other purpose | the given monitor will be installed on the stretcher for transporting seriously ill patients who require permanent monitoring. It must be mobile, with batteries, robust. | required
7 | Level of use | District hospital, provincial hospital, national hospital, specialized hospital.
8 | Clinical department/ward | Intensive Care Unit, Emergency, Recovery, Medical Imaging, Outpatient Services, Inpatient Care, Specialized Diagnostics, Specialized Treatment
9 | Overwiew of functional requirements | Patient vital signs monitor, 6 physiological parameters; ECG, heart rate, respiration rate, Non-invasive blood pressure and pulseoximetry and temperature. Suitable for adult, paediatric, AC and battery powered, including accessories. | required
10 | General configuration | Monitoring 6-parameters: ECG and Heart Rate (HR),Respiratory Rate (RR), SpO₂, non-invasive bloodpressure, and Temperature. | required
|  | Colour flat panel display size 10-12 inches. | required
|  | Multi-waveform and parameters visualization,showing minim 6 waveforms simultaneously | required
|  | Possibility to remove unwanted parameters from display | required
|  | Defibrillator syncronisation and protection | required
|  | Defibrillation recovery time | ≤10 sec
|  | Possibility to adjuste signal amplitude and sensitivity | required
|  | Data and network interface, LAN, USB or equivalent | required
|  | Suitable for standard bed/wall rail and pole standmount. | required
|  | Robust design for use in demanding environment | required
|  | Designed for frequent and easy dismount anddisinfection with hospital-grade products | required
|  | Automatic self-test and continuous system monitoring | required
|  | Built-in rechargeable battery, autonomy of minim 2 hours | required
|  | Must have alarm information displayed | required
|  | Visual and audible alarm, recall of alarm events | required
|  | Real-time monitoring of battery capacity, when the battery power is insufficient, low battery voltage alarm indication will display on LCD screen | required
TECHNICAL CHARACTERISTICS
11 | DISPLAY
|  | 10-12 inches, color | required
|  | resolution | minimum 800x680
|  | display multiple traces of signal waveforms for ECG | required
|  | showing waveforms simultaneously | minim 6
|  | Real-time show of battery capacity | required
|  | configure the display color, font and waveforms layout for user customization | required
ELECTROCARDIOGRAM (ECG) AND HEART RATE (HR), PR
|  | Input dynamic range | minim range ±(0.5mVp ~ 5mVp)
|  | ECG is derived from a minimum of 5-leads I, II, III,aVR, aVL, aVF, V. | required
|  | Accuracy of the HR is equal or better | ±1%
|  | HR measurement range | 15 – 300 bpm
|  | Arrhythmia analysis on the ECG | required
|  | Pacemaker detection | required
|  | sensitivity | 2.5 - 40 mm/mV
|  | RESP rate measuring range (RR) | 6 - 120 rpm
|  | RESP rate accuracy | max. ±5%
|  | speed 6.25mm/s, 12.5mm/s, 25mm/s, 50mm/s | required
|  | Differential input impedance | ≥10MΩ
12 | OXYGEN SATURATION (SPO₂)
|  | measurement range | 5 - 100%
|  | resolution of the SpO₂ | 1% and better
|  | The accuracy of the SpO₂, between 70% - 100%. | ≤ 2 %
|  | HR detection range derived from the SpO₂ | 25 – 250 bpm and better
|  | The accuracy of the HR detection | max ± 2 bpm
|  | Setting alarm | required
|  | Alarm range: High: 1%~100%; Low: 0%~99% | required
13 | NON-INVASIVE BLOOD PRESSURE NIBP
|  | Measuring method | oscillometric
|  | measured modes | manual and automatic
|  | NIBP adult range | equal or better than: 10 – 270 mmHg.
|  | unit allows for adjustable inflation pressure | required
|  | Accuracy of pressure measurement | max ±3 mmHg
|  | Typical measurement time | <30 seconds
|  | Initial cuff inflation pressure | required
14 | TEMPERATURE (T)
|  | measurement range | equal to or better than: 20 - 50°C.
|  | temperature resolution | 0.1°C or better
|  | Measuring site | body surface
|  | TEMP alarm limit setting range | required
15 | PRINTER | integrated thermal Printer | required
|  | used to print patient information, measurement numerics, up to three waveforms | required
16 | accessories | Adult SpO2 Finger clip Sensor | required
|  | Pediatric SpO2 Finger clip sensor | required
|  | Adult NIBP cuff | required
|  | Middle-sized Pediatric NIBP Cuff (10cm~19cm) | required
|  | Body TEMP probe | required
|  | ECG cable | required
|  | ECG electrode | required
|  | Power cord | required
|  | Printing paper | 10 rolls
UTILITY REQUIREMENTS
17 | Electrical | The voltage of the electric current | 220, ±10%; 50 Hz
|  | Plug type in compliance with required local/national standards and regulations | required
|  | Automatic switch to battery in case of power failure,automatic recharge on connection to mains. | required
|  | Electrical protection provided by fuses in both live andneutral supply lines | required
|  | Main cable at least 2,5 meter | required
|  | autonomy minim 2hours | required
TRAINING, INSTALLATION AND UTILISATION
18 | Transportation | Supplier must include the transportation to the final healthcare facility | required
| Installation | Supplier must perform the installation, safety and operation checks before handover | required
| Training of user/s | Training of users in hardware and software (including applications): | required
|  | • Basic operation of the system | required
|  | • Advanced operation of the system | required
|  | • Maintenance routine | required
WARRANTY AND MAINTENANCE
19 | Warranty | minim 24 months | required
| Preventive maintenance | Included in the warranty according with the interval and period recommended by the manufacturer. All nececary consumables and materials needed for maintenance during the given period to be included in the offer. | required
| Software / Hardware upgrade availability | (any avaiable updated during the warranty period) Included | required
DOCUMENTATION
20 | Documentation requirements | All supporting documentation, Technical/Maintenance manuals to be supplied in English. User manual in romanian language, according with LAW No.102 from 09-06-2017 regarding medical devices including: | required
|  | • Manual procedures required for calibration and routine maintenance | required
|  | • Troubleshooting protocols | required
|  | • Cleaning and disinfection protocols | required
|  | Contact details of manufacturer, supplier and local service agent to be provided | required
SAFETY AND STANDARDS
21 | Regulatory Approval / Certification | CE certificate | required
| Standards for the manufacturer | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes Classified either under EU MDD 93/42/ECC, or under EUMDR 2017/745 as Class IIb device ISO 9001 Quality managementsystems IEC 60601- Medical electrical equipment IEC 80601-2-49:2018 Medical electrical equipment - Part 2-49: Particular requirements for the basic safety andessential performance of multifunction patientmonitors. | required

## surgical aspirator techspecs
SUCTION PUMP, (surgical aspirator)
NAME, CATEGORY AND CODING
Parameters |  |  | Minimum expected Specification | Proposed Specificacion (to be completed by the bidder) | Reference fom manufacturer original document / brochure / page where provided information may be verified by the evaluation committee
PURPOSE OF USE
1 | Clinical or other purpose | Device for removing liquids or gases by suction | required
2 | Level of use | District hospital, provincial hospital, national hospital, specialized hospital.
3 | Clinical department/ward | Intensive Care Unit, Emergency, Recovery, Medical Imaging, Outpatient Services, Inpatient Care, Specialized Diagnostics, Specialized Treatment
4 | Overwiew of functional requirements | Suction Pump designed to be used by medical facilities, hospitals, medical practices and clinics with its ability to be the optimal medical suction pump for all suction needs.
TECHNICAL CHARACTERISTICS
5 | General configuration
|  | With Stand, 4 antistatic castors, minimum 2 with breaks. | required
|  | Maintenance Free | required
|  | Continuous use | required
6 | suction pump | Vacuum capacity: 60 l/min ± 10% | required
|  | Adjustable vacuum, maximum 700 mbar | required
|  | Oilless pump | required
|  | Sound Level: maximum 50 dBA | required
7 | General configuration | integrated Vacuometer, Vacuometer control 0-1000 mbar | required
|  | Integrated vacuometer, vacuometer control 0-1000 mbar | required
|  | Vacuum setting | required
|  | Device body - plastic | required
|  | Vacuum entrance - double | required
|  | Left-right jar change - with electronic button | required
|  | Foot pedal | required
|  | Stand body - anti-static painted metal | required
|  | Aspirating tube: min 4 m, inner ø 7-9 mm, silicone, autoclavable | required
|  | 4 plastic jars, 4 l each, graduated, with plug, autoclavable (autoclave at 121°c) | required
8 | SAFETY
|  | antibacterial filter | required
|  | Automatic overflow switch device at the level of the jar's plug | required
|  | minim 1 security bottle, 300-400 ml, with overflow switch | required
9 | Other accessories
|  | patient Single use bacterial Filter , pcs, min. | 50
|  | aspiration unit filter(machine filter), (Replace filters as per the manufacturer's recommended schedule, for minim 2 years , NOTE- some manufacturer recommend change outlet filter every 2 monts, some 1/2year, looks dirty or every accidental suction of liquid into pump) | full set during warranty period
UTILITY REQUIREMENTS
10 | Electrical and mecanical | The voltage of the electric current | 220, ±10%; 50 Hz
|  | Plug type in compliance with required local/national standards and regulations | required
|  | Appropriate external device to protect the equipment against over-voltage and over-current line conditions (between plug and socket) | required
TRAINING, INSTALLATION AND UTILISATION
11 | Transportation | Supplier must include the transportation to the final healthcare facility | required
| Installation | Supplier must perform the installation, safety and operation checks before handover | required
| Training of user/s | Training of users : | required
|  | • Basic operation of the pump | required
|  | • Advanced operation of the pump | required
|  | • Maintenance routine | required
WARRANTY AND MAINTENANCE
12 | Warranty | minim 24 months | required
| Preventive maintenance | Included in the warranty according with the interval and period recommended by the manufacturer. All nececary consumables and materials needed for maintenance during the given period to be included in the offer. | required
DOCUMENTATION
13 | Documentation requirements | All supporting documentation, Technical/Maintenance manuals to be supplied in English. User manual in romanian language, according with LAW No.102 from 09-06-2017 regarding medical devices including: | required
|  | • Manual procedures required for calibration and routine maintenance | required
|  | • Troubleshooting protocols | required
|  | • Cleaning and disinfection protocols | required
SAFETY AND STANDARDS
14 | Regulatory Approval / Certification | EC declaration of conformity. | required
| Standards for the manufacturer | ISO 13485:2003 Medical devices -- Quality management systems ISO 9001 Quality managementsystems IEC 60601- Medical electrical equipment ISO 10079-1 Medical suction equipment -- Part 1: Electrically powered suction equipment | required