## INR POCT
Annex 1
ITB 2024/EURO/MDA/0013
Procurement of the INR_POCT and Vein detector devices for the STROKE centers within the country hospitals
List of requirements and specifications
NAME, CATEGORY AND CODING
1 | Quantity | 13 |  |  |  | Bidder name: |  |  | Bidder name:
2 | Generic name | Portable PT/INR Point-of-Care Testing (POCT) device |  |  |  | Manufacturer |  |  | Manufacturer
3 | Level of use | District hospital, raion hospital, national hospital, specialized hospital. |  |  |  | Model |  |  | Model
4 | Place of use | Intensive Care Unit, Emergency, Recovery, Outpatient Services, Inpatient Care. |  |  |  | Country of origin: |  |  | Country of origin:
PURPOSE OF USE
1 | Clinical or other purpose | The handheld analyzer performs coagulation tests at the point of care (POC), meaning it can be used anywhere in the hospital near the patient. This allows for near-immediate results, significantly reducing the turnaround time (TAT) compared to traditional laboratory testing. |  |  |  | Meet WHO SPECS | NOTE DEVIATION |  | Meet WHO SPECS | NOTE DEVIATION
2 | Overview of functional requirements | Analyzers that perform Prothrombin Time (PT) assay and International Normalized Ratio (INR) testing through strips and/or cartridges
Yes
No
Item | Group | Category | Description | Acceptability Criteria | Meet WHO SPECS | Proposed Specificacion (to be completed by the bidder) | Reference document / brochure / page | Meet WHO SPECS | Proposed Specificacion (to be completed by the bidder) | Reference document / brochure / page | Unknown
1 | General Technical Characteristics | Type | Portable/Handheld | Mandatory
2 | General Technical Characteristics | Detection method | Electrochemical, Immunofluorescence, Microfluidics or other, good to performs coagulation tests | Mandatory
3 | General Technical Characteristics | Analysis time | <2 min | Mandatory
4 | General Technical Characteristics | Reagent type | Strips, cartridges or similar technology | Mandatory
5 | General Technical Characteristics | Calibration | Autotest on startup | Mandatory
6 | General Technical Characteristics | Memory | >200 test results | Mandatory
7 | Detailed Technical Characteristics | Test Menu | International Normalized Ratio (INR) | Mandatory
8 | Detailed Technical Characteristics | International Normalized Ratio (INR) range | 1.0 - 7.5 or better | Mandatory
| Detailed Technical Characteristics | Precision, accuracy of tests for Prothrombin | 3.5% and better
9 | Detailed Technical Characteristics | Sample volume | Less than 10 ul | Mandatory
10 | Detailed Technical Characteristics | Alarms | Low battery | Mandatory
11 | Detailed Technical Characteristics | Data and software management | Should have communication port or WIFI connection for data management | Mandatory
12 | Accessories | Carrying case | 1 pc | Mandatory
13 | Accessories | Battery charger | 1 pc | Mandatory
14 | Consumables | Strips/cartridges or similar | enough for 2000 tests | Mandatory
15 | Electrical | Working continous time (one charge) | not less than 4 hours | Mandatory
16 | Power requirement | 220-240V / 50Hz |  | Mandatory
17 | Safety and Standards | Regulations | (EU) 2017/746 or 98/79/EC or DoC | Mandatory
18 | Safety and Standards | International standards | ISO 13485 | Mandatory
19 | Safety and Standards | International standards | ISO 14001:2015 Environmental management systems — Requirements with guidance for use | Mandatory
20 | Safety and Standards | Local Regulations | Registration on Ministry of Health of the Republic of Moldova / Agency of Medicines and Medical Devices | Mandatory
21 | Safety and Standards | Regulations | Manufacturer`s authorization | Mandatory
22 | Warranty and maintenance | Service Conditions | Period: at least 24 months | Mandatory
23 | Warranty and maintenance | Service Conditions | Maximum response time: 72 hours | Mandatory
24 | Warranty and maintenance | Service Conditions | Preventive maintenance tasks included in the warranty twice a year | Mandatory
25 | Installation |  | Supplier must perform the safety and operational checks before handover | Mandatory
26 | Transportation |  | Supplier must include the transportation to the final destination | Mandatory
27 | Training and utilisation |  | Training for at least 3 medical staff who will use the item at each site at the beneficiary's location | Mandatory
28 | Training and utilisation |  | Training in Romanian or Russian language (if not compliant specify) | Mandatory
29 | Training and utilisation |  | Training on basic and advanced features and functions | Mandatory
30 | Training and utilisation |  | Training on preventive maintenance, daily maintenance and cleaning procedures | Mandatory
31 | Training and utilisation |  | Training to identify and resolve common faults/problems | Mandatory
32 | Documentation |  | All supporting documentation must be presented at least in English and, when available, in Romanian or Russian | Mandatory
33 | Documentation |  | Manual procedures required for local calibration and routine maintenance | Mandatory
34 | Documentation |  | Troubleshooting, cleaning and disinfection protocols | Mandatory
35 | Documentation |  | List of spares and accessories, with their part numbers and cost. | Mandatory
36 | Documentation |  | Two user’s manuals, preferable in Romanian language (also in electronic format) | Mandatory
37 | Documentation |  | Two service manual, preferable in Romanian language (also in electronic format) | Mandatory
38 | Local representative/partner | To provide service for the product | Located in Moldova/Romania (specify address) | Mandatory
39 | Local representative/partner | To provide service for the product | Name | Mandatory
40 | Local representative/partner | To provide service for the product | Street / N° | Mandatory
41 | Local representative/partner | To provide service for the product | ZIP / City | Mandatory
42 | Local representative/partner | To provide service for the product | Country | Mandatory
43 | Local representative/partner | To provide service for the product | Contact person and phone number | Mandatory
44 | Local representative/partner | To provide service for the product | Years of experience in medical equipment procurement | Mandatory
45 | Local representative/partner | To provide service for the product | Number of technicians skilled for the offered equipment (specify) | Mandatory

## vein detector
Annex 1
RFQ 2024/EURO/MDA/0001
Procurement of the INR_POCT and Vein detector devices for the STROKE centers within the country hospitals
List of requirements and specifications
NAME, CATEGORY AND CODING
1 | Quantity | 13 |  |  |  | Bidder name: |  |  | Bidder name:
2 | Generic name | Portable PROFESSIONAL VEIN finder device |  |  |  | Manufacturer |  |  | Manufacturer
3 | Level of use | District hospital, raion hospital, national hospital, specialized hospital. |  |  |  | Model |  |  | Model
4 | Place of use | Intensive Care Unit, Emergency, Recovery, Outpatient Services, Inpatient Care. |  |  |  | Country of origin: |  |  | Country of origin:
PURPOSE OF USE
1 | Clinical or other purpose | Medical Vein Finder is a high-quality imaging device designed to help medical professionals easily locate veins in patients. Using infrared technology and high-resolution imaging, this device offers a clear and detailed view of veins that may be difficult to see with the naked eye. Applications Peripheral IV insertion Blood draws Catheter insertion Varicose vein treatment Dermatology Other medical procedures requiring vein access |  |  |  | Meet WHO SPECS | NOTE DEVIATION |  | Meet WHO SPECS | NOTE DEVIATION
2 | Overview of functional requirements | Vein finder operates by using infrared light to detect veins beneath the skin, then projecting the position of the veins on the skin surface directly above the veins. Qualified medical personnel can observe the vasculature as displayed to assist them in finding a vein of the right size and position for venipuncture and other medical procedures requiring the location of superficial veins. No special training is required to operate this device.
Yes
No
Item | Group | Category | Description | Acceptability Criteria | Meet WHO SPECS | Proposed Specificacion (to be completed by the bidder) | Reference document / brochure / page | Meet WHO SPECS | Proposed Specificacion (to be completed by the bidder) | Reference document / brochure / page | Unknown
1 | General Technical Characteristics | PROFESSIONAL VEIN FINDER, medical grade design | Infrared technology for enhanced vein visibility, non-contact | Mandatory
2 | General Technical Characteristics | Detection method | Infrared technology | Mandatory
3 | General Technical Characteristics | Analysis time | Real-time vein imaging and continuously operating | Mandatory
4 | General Technical Characteristics | Suitable | for adults, children and newborns | Mandatory
5 | General Technical Characteristics | Display | minim 2 ", multi color | Mandatory
6 | General Technical Characteristics | Brightness | minim 5 Levels of Brightness | Mandatory
7 | General Technical Characteristics | Mixture Colors | Show the depth information of veins | Mandatory
8 | Detailed Technical Characteristics | Software hair removal | Eliminate hair interference in vein image | Mandatory
9 | Detailed Technical Characteristics | Infrared Wave range | between 750-950nm | Mandatory
10 | Detailed Technical Characteristics | Accuracy of blood vessel position | better than ±0.3mm | Mandatory
11 | Detailed Technical Characteristics | Accuracy of blood vessel resolution | better than ±0.3mm | Mandatory
12 | Detailed Technical Characteristics | detection distance | range of 15-25 cm, -2+10 cm deviation | Mandatory
13 | Detailed Technical Characteristics | detection depth | better than in range 1-10mm | Mandatory
14 | Detailed Technical Characteristics | image save | possibility to save and export images to PC | Mandatory
15 | Detailed Technical Characteristics | Battery status | Must be displayed | Mandatory
16 | Detailed Technical Characteristics | Battery saver | Enter into low power consumption mode when the user needs short intervals and can be waked up quickly. Sleep time can be set. | Mandatory
17 | Detailed Technical Characteristics | Battery type | litium or better type, recharged type | Mandatory
18 | Detailed Technical Characteristics | Working time | minim 3 hours of continuosly operating | Mandatory
19 | Accessories | Carrying case | 1 pc | Mandatory
20 | Accessories | Battery charger | 1 pc | Mandatory
21 | Consumables | calibration | if device required periodical calibration, calibration accessories must be included in set | Mandatory
22 | Power requirement | power supply | 100V-240V 50Hz-60Hz | Mandatory
23 | Safety and Standards | International standards | CE mark Medical Device compliant with Regulation (EU) 2017/745 | Mandatory
24 | Safety and Standards | International standards | IEC 60601-1-2 | Mandatory
25 | Safety and Standards | Regulations | Manufacturer`s authorization | Mandatory
26 | Warranty and maintenance | Service Conditions | Period: at least 24 months | Mandatory
27 | Warranty and maintenance | Service Conditions | Maximum response time: 72 hours | Mandatory
28 | Installation |  | Supplier must perform the safety and operational checks before handover | Mandatory
29 | Transportation |  | Supplier must include the transportation to the final destination | Mandatory
30 | Training and utilisation |  | Training on basic and advanced features and functions | Mandatory
31 | Training and utilisation |  | Training on preventive maintenance, daily maintenance and cleaning procedures | Mandatory
32 | Training and utilisation |  | Training to identify and resolve common faults/problems | Mandatory
33 | Documentation |  | All supporting documentation must be presented at least in in Romanian and English | Mandatory
34 | Documentation |  | Manual procedures required for local calibration and routine maintenance if needed | Mandatory
35 | Documentation |  | Troubleshooting, cleaning and disinfection protocols | Mandatory
36 | Local representative/partner | To provide service for the product | Located in Moldova/Romania (specify address) | Mandatory
37 | Local representative/partner | To provide service for the product | Name | Mandatory
38 | Local representative/partner | To provide service for the product | Street / N° | Mandatory
39 | Local representative/partner | To provide service for the product | ZIP / City | Mandatory
40 | Local representative/partner | To provide service for the product | Country | Mandatory
41 | Local representative/partner | To provide service for the product | Contact person and phone number | Mandatory
42 | Local representative/partner | To provide service for the product | Years of experience in medical equipment procurement | Mandatory
43 | Local representative/partner | To provide service for the product | Number of technicians skilled for the offered equipment (specify) | Mandatory

## 16-Dispo.padeles ad.ped. (hidden)
Item Disposable padeles for adults and pediatricc