## blood refrigerator
Annex 1
ITB 2024/EURO/MDA/0022
Procurement of specialised blood storage refrigeration devices for the National Blood Transfusion Center and country hospitals
NAME, CATEGORY AND CODING
| 1 | Quantity | 20 |  |  |  | Bidder name:
| 2 | Generic name | Standard electric blood bank refrigerators |  |  |  | Manufacturer
| 3 | Level of use | Twente district hospitals, raion hospital, national hospital, specialized clinics, blood stations |  |  |  | Model
| 4 | Place of use | Diagnostic wards / departments |  |  |  | Country of origin:
PURPOSE OF USE
| 1 | Clinical or other purpose | A refrigerator for storing whole blood or red cell packs in a blood bank, 2~6℃ |  |  |  | Meet WHO SPECS | NOTE DEVIATION
| 2 | Overview of functional requirements | Compression type refrigerator that uses CFC-free refrigerant gas and electricity supply from the national grid
| Item | Group | Category | Description | Acceptability Criteria | Meet WHO SPECS | Proposed Specificacion (to be completed by the bidder) | Reference from official document / brochure / page
| 1 | Construction type | Body type and materials
| 2 |  |  | Vertical (Upright) | Mandatory
| 3 |  |  | Internal material Stainless steel (min. 22g) | Mandatory
| 4 |  |  | External material Stainless steel or Corrosion Resistant (CR at least 1mm thickness) | Mandatory
| 5 |  |  | CFC-free | Mandatory
| 6 |  |  | Drawers: Roll out type with perforation for air circulation | Mandatory
| 7 |  |  | Door: Single door, Glass or partial glass | Mandatory
| 8 |  |  | Door lock with key or password | Mandatory
| 9 |  | Drawers | 6 ±1pcs, Stainless steel, perforated for air recirculation | Mandatory
| 10 |  |  | Stainless steel drawers secured with septa, which will divide the product storage space (bags of blood components); | Mandatory
| 11 |  | Casters | 4（2 caster with brake) | Mandatory
| 12 |  | Cooling method | Forced Air cooling | Mandatory
| 13 |  |  | Forced Air cooling switches off automatically when open the door, ensures a uniform temperature and minimizes temperature deviations. | Mandatory
| 14 |  | Useful Capacity(L) | 360-400l +-3% | Mandatory
| 15 |  | Performance
| 16 |  | External Ambient Temperature: | Performs in an ambient temperature of +10 to +43 °C, or SN ~ T Climate Class | Mandatory
| 17 |  | Working Temperature range | 2~6℃ | Mandatory
| 18 |  | Defrost | automatic | Mandatory
| 19 |  |  | force defrost | Mandatory
| 20 |  | Constant Temperature Control | Temperature control precision of ±1℃ inside the cabinet | Mandatory
| 21 |  |  | Built-in sensors for upper/lower zone temperature, ambient temperature, evaporator temperature, operation control | Mandatory
| 22 |  |  | record and display parameters for 7 days minim | Mandatory
| 23 |  | Temperature controller | Microprocessor | Mandatory
| 24 |  | Display | Digital, or LCD, or TFT | Mandatory
| 25 | Compressor | Low-noise compressor | not more than 53dB noise | Mandatory
| 26 | Security System
| 27 |  | Alarm system | Audible and visible alarm | Mandatory
| 28 |  |  | High temperature | Mandatory
| 29 |  |  | Low temperature | Mandatory
| 30 |  |  | Sensor fault | Mandatory
| 31 |  |  | Door opening | Mandatory
| 32 |  |  | Power failure | Mandatory
| 33 |  |  | Low backup battery | Mandatory
| 34 |  |  | Controller failure | Mandatory
| 35 |  | Power supply | 220/240V 50Hz ±10%, euro plug with ground | Mandatory
| 36 |  |  | inside lighting system, LED | Mandatory
| 37 |  |  | Equipment meets electrical safety specifications such as that of IEC | Mandatory
| 38 |  |  |  | Mandatory
| 39 |  | Refrigerant | environmentally friendly, EC permitted and approved | Mandatory
| 40 |  |  |  | Mandatory
| 41 | Software | data | Data logging, and data transfer possibility (USB , Ethernet interface for data management) | Mandatory
| 42 |  |  |  | Mandatory
| 43 | Safety and Standards - Hardware | International Regulations | ISO 9001:2015 Quality management systems — Requirements | Mandatory
| 44 | Safety and Standards - Hardware | International Regulations | ISO 13485 and/or MDR (EU) 2017/745, Class IIA | Mandatory
| 45 | Warranty and maintenance | Service Conditions no restrictions | Warranty Period: at least 36 months | Mandatory
| 46 | installation and setting of workstations | included for full set | Supplier must provide a full operational and fully working working devices | Mandatory
| 47 | Installation |  | Supplier must perform the installation, safety and operation checks before handover | Mandatory
| 48 | Installation |  | Equipment assembly and commissioning on site at beneficiary location | Mandatory
| 49 | Transportation |  | Supplier must include the transportation to the final destination | Mandatory
| 50 | Documentation |  | User’s manuals, in Romanian language (also in electronic format) | Mandatory
| 51 | Documentation |  | Service manual, preferable in Romanian language or English (also in electronic form) | Mandatory
| 52 | Documentation |  | Installation and setup notes / project documentation of delivered state | Mandatory
| 53 | Local representative/partner | To provide service for the product | Located in Moldova | Mandatory
| 54 | Local representative/partner | To provide service for the product | Experience in medical equipment procurement and service | Mandatory
| 55 | Year of production |  | Must be manufactured not before 2024 year | Mandatory

## plasma freezers
Annex 1
ITB 2024/EURO/MDA/0022
Procurement of specialised blood storage refrigeration devices for the National Blood Transfusion Center and country hospitals
List of requirements and specifications
NAME, CATEGORY AND CODING
| 1 | Quantity | 20 |  |  |  | Bidder name:
| 2 | Generic name | Blood plasma freezers |  |  |  | Manufacturer
| 3 | Level of use | Twente district hospitals, raion hospital, national hospital, specialised clinics, blood stations |  |  |  | Model
| 4 | Place of use | Diagnostic wards / departments |  |  |  | Country of origin:
PURPOSE OF USE
| 1 | Clinical or other purpose | To freeze and store plasma in a blood bank, Operating temperature, –35 °C to –40 °C |  |  |  | Meet WHO SPECS | NOTE DEVIATION
| 2 | Overview of functional requirements | Compression freezer with CFC-free refrigerant gas and electricity supply from the national grid
| Item | Group | Category | Description | Acceptability Criteria | Meet WHO SPECS | Proposed Specificacion (to be completed by the bidder) | Reference from official document / brochure / page
| 1 | Construction type | Body type and materials
| 2 |  |  | Vertical (Upright) | Mandatory
| 3 |  |  | Internal material Stainless steel (304) | Mandatory
| 4 |  |  | External material Stainless steel or Corrosion Resistant (CR at least 1mm thickness), Color sprayed steel | Mandatory
| 5 |  |  | CFC-free | Mandatory
| 6 |  |  | Drawers: Roll out type with perforation for air circulation | Mandatory
| 7 |  |  | Door: Single door | Mandatory
| 8 |  |  | Door lock with key or password | Mandatory
| 9 |  | Drawers | 5 ±1pcs, perforated for air recirculation, must be height adjustable | Mandatory
| 10 |  |  | Equipped with special plasma bag clips | Removed
| 11 |  | Casters | 2+(2 leveling feet), or 4（2 caster with brake) | Mandatory
| 12 |  | Cooling method | Forced Air cooling, nofrost | Mandatory
| 13 |  |  | Automatic switch-off of the evaporator fans at each door opening to reduce heat transfer | Mandatory
| 14 |  | Useful Capacity(L) | 340-380l +-3% | Mandatory
| 15 |  | Performance
| 16 |  | External Ambient Temperature: | Performs in an ambient temperature of +10 to +43 °C, or SN ~ T Climate Class | Mandatory
| 17 |  | Working Temperature range | up -35℃ | Mandatory
|  |  | Cooling performance | -40℃ | Mandatory
| 18 |  | Defrost | manual | Mandatory
| 20 |  | Constant Temperature Control | Temperature control precision of ±1℃ inside the cabinet in range up -60C | Mandatory
| 21 |  |  | Built-in sensors for upper/lower zone temperature, ambient temperature, operation control | Mandatory
| 22 |  |  | record and display parameters for 7 days minim or min 50000 data logs | Mandatory
| 23 |  | Temperature controller | Microprocessor | Mandatory
| 24 |  | Display | Digital, or LCD, or TFT | Mandatory
| 25 | Compressor | Low-noise compressor | not more than 58dB noise | Mandatory
| 26 | Security System
| 27 |  | Alarm system | Audible and visible alarm | Mandatory
| 28 |  |  | High temperature | Mandatory
| 29 |  |  | Low temperature | Mandatory
| 30 |  |  | Sensor fault | Mandatory
| 31 |  |  | Door opening | Mandatory
| 32 |  |  | Power failure | Mandatory
| 33 |  |  | Low backup battery | Mandatory
|  |  |  | Condenser clean | Mandatory
| 34 |  |  | Controller failure | Mandatory
|  |  |  | RS485,Remote alarm contact | Optional
| 35 |  | Power supply | 220/240V 50Hz ±10%, euro plug with ground | Mandatory
|  |  |  | Extended autonomy in case of power failure | Mandatory
| 36 |  |  | inside lighting system, LED | Mandatory
| 37 |  |  | Equipment meets electrical safety specifications such as that of IEC | Mandatory
| 38 |  | Backup battery | Backup battery for 24 hours | Mandatory
| 39 |  | Refrigerant | environmentally friendly, EC permitted and approved | Mandatory
| 40 |  |  |  | Mandatory
| 41 | Software | data | Data logging, and data transfer possibility (USB , Ethernet interface for data management) | Mandatory
| 42 |  |  |  | Mandatory
| 43 | Safety and Standards - Hardware | International Regulations | ISO 9001:2015 Quality management systems — Requirements | Mandatory
| 44 | Safety and Standards - Hardware | International Regulations | ISO 13485 and/or MDR (EU) 2017/745, Class IIA | Mandatory
| 45 | Warranty and maintenance | Full Service Conditions no restrictions | Warranty Period: min 36 months | Mandatory
| 46 | installation and setting of workstations | included for full set | Supplier must provide a full operational and fully working working devices | Mandatory
| 47 | Installation |  | Supplier must perform the installation, safety and operation checks before handover | Mandatory
| 48 | Installation |  | Equipment assembly and commissioning on site at beneficiary location | Mandatory
| 49 | Transportation |  | Supplier must include the transportation to the final destination | Mandatory
| 50 | Documentation |  | User’s manuals, in Romanian language (also in electronic format) | Mandatory
| 51 | Documentation |  | Service manual, preferable in Romanian language or English (also in electronic form) | Mandatory
| 52 | Documentation |  | Installation and setup notes / project documentation of delivered state | Mandatory
| 53 | Local representative/partner | To provide service for the product | Located in Moldova | Mandatory
| 54 | Local representative/partner | To provide service for the product | Experience in medical equipment procurement and service | Mandatory
| 55 | Year of production |  | Must be manufactured not before 2024 year | Mandatory

## 16-Dispo.padeles ad.ped. (hidden)
Item Disposable padeles for adults and pediatricc