## 1 list equipmnet
Annex 1
ITB 2025/EURO/MDA/0002
List of equipment
Annex A | Ultrasonic bone scalpel (bisturiu cu ultrasunete pentru oase)
Annex B | System for ablation of bone tumors by radiofrequency (Sistem pentru ablatia a tumorilor osului prin radiofrecventa)
Annex C | Advanced dental unit (Unitate stomatologica caracteristici avansate)
Annex D | Interventional Pulmonology workstation consisting of, electrosurgery, argon-plasma coagulation, cryosurgery CO2. (Statie dedicata bronhoscopiei interventional e formata din Unitatea chirurgicala, Unitatea de coagulare cu Argon-Plasma, Unitatea Criochirurgie cu CO2)
Annex E | Electrosurgical unit with argon-plasma dedicated for gastrointestinal procedures such as ESD (Electrocoagulator predestinat pentru chirurgia mucoasei tractului digestiv si cailor biliare)
Annex F | Ultrasonic bath for surgical instrument 30L (Baie cu ultrasunet pentru instrumente chirurgicale 30 de litri)
Annex G | Aplicator Skin Cancer Treatment with Brachytherapy (BRA VOS). Aplicator Brahiterapie tumori cutanate
Annex H | Ultrasonic surgery device type Harmonic, for thoracoscopic interventions (generator de ultrasunete cu coagulare tip harmonic

## A-ultrasonic bone scalpel
Annex A
ITB 2025/EURO/MDA/0002
Procurement of the Ultrasonic bone scalpel
List of requirements and specifications
NAME, CATEGORY AND CODING
1 | Quantity | 1 |  |  |  | Bidder name:
2 | Generic name | Ultrasonic bone scalpel |  |  |  | Manufacturer
3 | Level of use | National hospital/ specialized clinics |  |  |  | Model
4 | Place of use | Institute of Oncology |  |  |  | Country of origin:
PURPOSE OF USE
1 | Clinical or other purpose | Ultrasonic bone scalpel intended for use in selective fragmentation, resection, emulsification, irrigation of hard tissue (bone, calcified lesions) in orthopedic and spinal oncological surgery. |  |  |  | Meet WHO SPECS | NOTE DEVIATION
2 | Overview of functional requirements | Function: used for fine orthopedic surgery and spinal oncological surgery . Orthopedics with functions such as cutting, grinding and drilling of bones.
Item | Group | Category | Description | Acceptability Criteria | Meet WHO SPECS | Proposed Specificacion (to be completed by the bidder) | Reference from official document / brochure / page
1 | Construction type | Body type and materials
2 | Used for |  |  | Mandatory
3 | Classifications | Main unit | Control console (generator) | Mandatory
4 |  | Interface | screen | Mandatory
5 |  | parameters | Possibility to modify output parameters | Mandatory
6 |  | Working frequency | ≥20kHz | Mandatory
7 |  | Working frequency | Automatic frequency tracking function to ensure the best power output effect in real time | Mandatory
8 |  | Power adjustment | 0-100% adjustable | Mandatory
9 |  | Self-checking system | Ultrasonic bone scalpel has a fast self-checking function, which can alarm the faulty handle and cutter head | Mandatory
10 |  | Handpiece (ultrasonic scalpel) | Yes, 2 pcs | Mandatory
11 |  | Ultrasonic scalpel connection cable | minimum 2 meters | Mandatory
12 |  | Protection and safety | Protects against destruction of soft tissue structures during contact | Mandatory
13 |  | Control | Foot pedal | Mandatory
14 |  | Control panel | The Ultrasonic bone scalpel is displayed in a color LCD | Mandatory
15 |  |  | Irrigation system | Mandatory
16 |  |  | Irrigation solution fixing bracket | Mandatory
17 |  | Trolley (stand) | Minimum 4 wheels, minimum one with brake, minimum one shelf | Mandatory
18 |  | Accesssories | Set of scalpel tips for scalpel, sterile, single use, 10 pieces per set, different sizes (e.g. 10 mm, 20mm), and other fixing parts | Mandatory
19 |  | Set of scalpel tips for scalpel, sterile, single use, 10 pieces per set, different sizes (e.g. 10 mm, 20mm), and other fixing parts | 2 sets of different sizes | Mandatory
20 |  | Scalpel tips for scalpel, reusable, (Note: only if they are manufacturing) | 2 pcs | Optional
21 |  | Tips of different sizes and configurations | ≥5 pcs | Mandatory
22 |  | Irrigation system tubing set | 20 sets | Mandatory
23 |  | Flushing and suction pipelines | ≥3m | Mandatory
24 |  | Accessories | Accessories required for full use of the ultrasonic scalpel | Mandatory
25 |  | Maintenance | Scalpel maintenance accessories (as appropriate according to manufacturer's recommendations: cleaning set, assembly keys, etc.) | Mandatory
26 | POWER SUPPLY | EU plug with ground | 220-240 [V] +/- 10% 50/60 [Hz] | Mandatory
27 | Safety and Standards | International Regulations | ISO 9001:2015 Quality management systems — Requirements | Mandatory
28 |  | International Regulations | ISO 13485 | Mandatory
29 |  | International Regulations | CE marking or EU declaration of conformity | Mandatory
30 | Safety and Standards - Hardware | International Regulations | minimum certificates - IEC 60601-1 E96 General requirements for basic safety and essential performance. - IEC 60601-2-2 Medical Electrical Equipment – Part 2-2: Particular requirements for the safety of high frequency surgical equipment | Mandatory
31 | Warranty and maintenance | Service Conditions- no restrictions | Warranty Period: minim 2 years | Mandatory
32 | installation and setting of workstations | included for full set | Supplier must provide a full operational and fully working working devices | Mandatory
33 | Installation | included for full set | The Ultrasonic bone scalpel must include all accessories for normal operation, Supplier must perform the installation, safety and operation checks before handover. Training of user and techicians must be provided. Equipment assembly and commissioning on site at beneficiary location | Mandatory
34 | Transportation | included for full set | Supplier must include the transportation to the final destination | Mandatory
35 | Documentation | included for full set | User’s manuals, in Romanian language (also in electronic format) | Mandatory
36 | Documentation | included for full set | Service manual, preferable in Romanian language or English (also in electronic form) | Mandatory
37 | Documentation | included for full set | Installation and setup notes / project documentation of delivered state | Mandatory
38 | Local representative/partner | To provide service for the product | Located in Moldova | Mandatory
39 | Local representative/partner | To provide service for the product | Experience in medical equipment procurement and service of medical equipment, certificated from manufacturer with presented confirmation | Mandatory
40 | Year of production |  | Must be manufactured not before 2024 year | Mandatory

## B- ablation system
Annex B
ITB 2025/EURO/MDA/0002
Procurement of the Radiofrequency ablation unit
List of requirements and specifications
NAME, CATEGORY AND CODING
1 | Quantity | 1 |  |  |  | Bidder name:
2 | Generic name | An electrically-powered, noninvasive device designed to generate radio-frequency (RF) electrical current used to create heat via a percutaneously (through the skin) introduced electrode(s) in a precise location, at a controlled temperature, for focal ablation of non-cardiac tissues (e.g., non-cardiac nerves, tumours, precancerous tissue); it may be intended for coagulation however it is not intended for electrosurgical cutting. The generator connects via a delivery cable to an electrosurgical ablation probe/catheter (not included) to transmit the RF electrical current to the operative site; it might additionally be intended to heat an electrode tip (electrocautery). |  |  |  | Manufacturer
3 | Level of use | National hospital/ specialized clinics |  |  |  | Model
4 | Place of use | Institute of Oncology |  |  |  | Country of origin:
PURPOSE OF USE
1 | Clinical or other purpose | Radiofrequency ablation (RFA) of bone tumors is a minimally invasive medical treatment that uses radiofrequency energy to generate heat and destroy cancerous cells in bone tumors. |  |  |  | Meet WHO SPECS | NOTE DEVIATION
2 | Overview of functional requirements
Item | Group | Category | Description | Acceptability Criteria | Meet WHO SPECS | Proposed Specificacion (to be completed by the bidder) | Reference from official document / brochure / page
1 | Construction type | Body type and materials
2 | Used for |  | Minimally invasive applications, Palliative treatment by ablation of metastatic malignant lesions involving the bone, including the vertebral body , Ablation of benign bone tumors, such as osteoid osteoma | Mandatory
3 | Classifications
4 | TECHNICAL CHARACTERISTICS and PHYSICAL CHARACTERISTICS
5 |  | Energy Type | Radiofrequency | Mandatory
6 |  | Mechanism of Action | Thermal coagulation necrosis | Optional
7 |  | Energy Delivery Control Mechanism | Temperature controlled | Mandatory
8 |  | Probe Active Tip Size | 7mm, 10mm, 15mm, 20mm | Mandatory
9 |  | Number of independent thermocouples | minim 2 | Mandatory
10 |  | Number of probes | minim 2 | Mandatory
11 |  | Radiofrequency Generator Power Output Channels | minim 2 | Mandatory
12 |  | Radiofrequency Generator Power | minim 40 Watts total | Mandatory
13 |  | Output Power | minim 20 W/channel | Mandatory
14 |  | Temperature monitoring | The temperature measured by independent sensors/thermocouples in the ablation area should be displayed on the screen in real time | Mandatory
15 |  | Temperature monitoring | Monitors temperature and power output to reduce the risk of thermal damage to adjacent structures | Mandatory
16 |  | Probes | Recognizes probe size and automatically presets ablation time | Mandatory
17 |  | Probes | Allows the use of 2 probes simultaneously | Mandatory
18 |  | Channel control | Independent | Mandatory
19 |  | Channel control | ability to stop and start channels individually | Mandatory
20 |  | Adjusting parameters per probes | time | Mandatory
21 |  | Adjusting parameters per probes | temperature | Mandatory
22 |  | Adjusting parameters per probes | power | Mandatory
23 |  | Adjusting parameters per probes | rate | Mandatory
24 |  | Adjusting parameters per probes | impedance | Mandatory
25 |  | Pumping system | peristaltic pumps or similar system | Mandatory
27 |  | Pumping system | minim 2 (two) independent pumps | Mandatory
28 |  | Pumping system | included Connection cables | Mandatory
29 | ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
30 |  | STAND (trolley) | stand with at least 4 wheels, 2 with brakes, for fixing generator, pumps, accessories, depending on the system | Mandatory
31 |  | Bone ablation set consisting of | bipolar radiofrequency probe with internal cooling system, active working tip 20 mm | 10 sets
|  |  | sterile container and connecting tubes | 10 sets
32 |  | Double bone ablation set consisting of | bipolar radiofrequency probes with internal cooling system, active working tip 20 mm | 1 set
|  |  | 2 sterile containers and connecting tubes | 1 set
33 |  | Transpedicular vertebral bone access set | size 12G -13G | 1 set
|  |  | manual drill with depth marks | 1 set
|  |  | stiletto | 1 set
|  |  | cannula | 1 set
33.661 | POWER SUPPLY | EU plug with ground | 220-240 [V] +/- 10% 50/60 [Hz] | Mandatory
35 | Safety and Standards - Hardware | International Regulations | ISO 9001:2015 Quality management systems — Requirements | Mandatory
| Safety and Standards - Hardware | International Regulations | ISO 13485 Medical devices and/or MDR (EU) 2017/745, | Mandatory
| Safety and Standards - Hardware | International Regulations | Biocompatibility of patient contact components according to ISO 10993-1 | Mandatory
36 | Safety and Standards - Hardware | International Regulations | Safety Standard IEC 61010 Safety requirements for medical devices | Mandatory
| Safety and Standards - Hardware | International Regulations | CE marking or EU declaration of conformity | Mandatory
37 | Warranty and maintenance | Service Conditions- no restrictions | Warranty Period: minim 2 years | Mandatory
38 | installation and setting of workstations | included for full set | The Radiofrequency ablation must include all accessories for normal operation, Supplier must perform the installation, safety and operation checks before handover. Training of user and techicians must be provided. Equipment assembly and commissioning on site at beneficiary location. Supplier must provide a full operational and fully working working devices | Mandatory
39 | Transportation | included for full set | Supplier must include the transportation to the final destination | Mandatory
40 | Documentation | included for full set | User’s manuals, in Romanian language (also in electronic format) | Mandatory
41 | Documentation | included for full set | Service manual, preferable in Romanian language or English (also in electronic form) | Mandatory
42 | Documentation | included for full set | Installation and setup notes / project documentation of delivered state | Mandatory
43 | Local representative/partner | To provide service for the product | Located in Moldova | Mandatory
44 | Local representative/partner | To provide service for the product | Experience in medical equipment procurement and service of medical equipment, certificated from manufacturer with presented confirmation | Mandatory
45 | Year of production |  | Must be manufactured not before 2024 year | Mandatory

## C- stomatologic unit
Annex C
ITB 2025/EURO/MDA/0002
Procurement of the Stomatologic unit
List of requirements and specifications
NAME, CATEGORY AND CODING
1 | Quantity | 1 |  |  |  | Bidder name:
2 | Generic name | Dental Unit |  |  |  | Manufacturer
3 | Level of use | National hospital/ specialized clinics |  |  |  | Model
4 | Place of use | Institute of Oncology |  |  |  | Country of origin:
PURPOSE OF USE
1 | Clinical or other purpose | Examination/treatment chairs designed to facilitate dental examination, treatment, and/or minor surgical procedures. These chairs are typically adjustable up to a height that allows the healthcare staff to perform procedures while standing; the chairs usually include head and armrests, a reclining back that may be tilted from a vertical to a horizontal or near-horizontal position, and rotating capabilities to facilitate examination and/or treatment. |  |  |  | Meet WHO SPECS | NOTE DEVIATION
2 | Overview of functional requirements
Item | Group | Category | Description | Acceptability Criteria | Meet WHO SPECS | Proposed Specificacion (to be completed by the bidder) | Reference from official document / brochure / page
1 | Construction type | COLOR of dental unit | The color of all unit must be from cold clolour palette side (black, grey, green, blue, any shades) | Mandatory
2 | Dental Chair | Armrests | Armrests play a vital role in ensuring patient stability and relaxation. They provide a resting place for the patient's arms, allowing them to remain comfortable throughout the procedure. Armrests are essential for supporting patient stability and relaxation. They give the patient's arms a place to rest, allowing them to be comfortable throughout the treatment. Left and right armrest in set. | Mandatory
3 |  | Headrest | The headrest is a crucial part of a dental chair, as it supports the patient's head and neck. It can be adjusted to accommodate patients of different heights and positions. A properly positioned headrest improves patient comfort while also providing a sturdy platform for dental operations, allowing the dentist to work with more precision. Included | Mandatory
4 |  | Seat and backrest | The seat and backrest of a dental chair are designed to maximize patient comfort during treatment. They are often padded with high-quality cushioning materials, offering optimal support for extended periods. The ergonomic design assists patients in maintaining a comfortable position, decreasing strain and pain. Included | Mandatory
5 |  | Tilt angle | from 100° to 180° | Mandatory
6 |  | headrest | adjustable | Mandatory
7 |  | Ascent/descent | ≥ 40cm | Mandatory
8 |  | guaranteed weight capacity | ≥ 150kg | Mandatory
9 |  | Surface | resistant to chemical processing | Mandatory
10 |  | Control | pedal (wired or wireless) | Mandatory
11 |  | Positioning programs | minimum 3 | Mandatory
12 |  | Spittoon Bowl | bowl handle is removable and autoclavable | Mandatory
13 | Dentist chair | Surface | resistant to chemical processing | Mandatory
14 |  | Maximum load | ≥ 150kg | Mandatory
15 |  | Height | adjustable | Mandatory
16 |  | Rear support | adjustable | Mandatory
17 |  | Footrest | dotated with foot rest place | Mandatory
18 | Dentist's table | Turbine | included | Mandatory
19 |  | 3-way syringe | water, air, spray | Mandatory
20 |  | Scaler | Scaler with at least 6 tips with LED lighting | Mandatory
21 |  | Electric micromotor | included | Mandatory
22 |  | CONTRA-ANGLE piece | included | Mandatory
23 |  | control | Touch display | Mandatory
24 | ASSISTANT’S MODULE (arm) | Saliva aspirator | included | Mandatory
25 |  | Control panel | Control panel with the same functions as on the dentist's table | Mandatory
26 |  | Intraoral camera | the specifications at media description | Mandatory
27 |  | Photopolymer lamp | head that can rotate, at least 3 modes of operation suitable for different treatment needs | Mandatory
28 |  | Photopolymer lamp | light intensity - not less than 2000 mW/cm, adjustable | Mandatory
29 | Hidraulic module | ceramic or glass bowl | Removable ceramic or glass bowl | Mandatory
30 |  | Dirt filter | dotate with removable filter | Mandatory
31 |  | water volume | Possibility of programming the water volume | Mandatory
32 | Dental turbine | changing the bit | one thouch Button system for fast changing the bit | Mandatory
33 |  | Working pressure | from 2.5 to 2.7 bar +-10% | Mandatory
34 |  | light | doated with Light source | Mandatory
35 |  | holes | Applicable to the ISO international standard holes | Mandatory
36 |  | Noise | less 68dB | Mandatory
37 |  | RPM | up 300000 rot/min | Mandatory
38 | Dental Light source | Dental Light | The dental or operator light is used to illuminate inside the mouth or oral cavity of the patient seated on the dental chair | Mandatory
39 |  | Dental Light | pre-set programmes and a sterilisable guide | Mandatory
40 |  | LIGHT | LED light with minim 3 intensity with 3 axis movement | Mandatory
41 |  | power | Intensity is up to - 40000 Lux, adjustable | Mandatory
42 |  | colour | WARM, NEUTRAL AND COOL COLOUR | Mandatory
43 | Straight handpiece | RPM | ≥40000 rpm | Mandatory
44 |  | Adaptable to micromotor | fully adapted | Mandatory
45 |  | autoclave | temperature range of 121-134 degrees | Mandatory
46 |  | With light | light included | Mandatory
47 |  | Changing the cutter | button | Mandatory
48 |  | Water/air cooling | holes according ISO | Mandatory
49 | Contra-angle handpiece | RPM | ≥40000 rpm | Mandatory
50 |  | Adaptable to micromotor | fully adapted | Mandatory
51 |  | autoclave | temperature range of 121-134 degrees | Mandatory
52 |  | With light | light included | Mandatory
53 |  | Changing the cutter | button | Mandatory
54 |  | Water/air cooling | holes according ISO | Mandatory
55 | Scaler | 5 tips with LED lighting | included | Mandatory
56 |  | Detachable, autoclavable handle | temperature range of 121-134 degrees | Mandatory
57 |  | Working frequency | 25 to 30 KHz | Mandatory
58 |  | power | adjustable | Mandatory
59 |  | Tips type included in set | more than 4 types | Mandatory
60 |  | Instrument Tray | the instrument tray is a pivotal part of the dental chair setup, holding sterilized instruments and tools required for specific procedures | Mandatory
61 |  | Foot Control | Multifunctional Foot Pedal. The pedals have self-lock function and equipped with control of rise and down, backrest forward and backward, blowing, flushing .. | Mandatory
62 |  | Suction Unit | The suction unit is responsible for removing excess fluids, debris, and saliva from the patient's mouth. It helps maintain a clean and dry working area for the dentist. A dependable suction device aids in infection control increases vision, and improves patient comfort by minimizing excessive fluid pooling | Mandatory
63 | compressor | Medical air compressor | Oil-free | Mandatory
64 |  |  | enough air production power to operate the dental unit | Mandatory
65 | LIGHT | LIGHT | LED light with minim 3 intensity with 3 axis movement | Mandatory
|  | power | Intensity is up to - 40000 Lux, adjustable | Mandatory
|  | colour | WARM, NEUTRAL AND COOL COLOUR | Mandatory
66 | media | LCD monitor | minim 19” monitor certified as per the EEC Medical Devices Directive | Mandatory
67 |  | Intraoral Camera | multimedia system consists of an inbuilt camera with HD sensor, CCD type, and extensive depth of field. Depth of field with minimum range from 5 to 70 mm. | Mandatory
68 |  | Intraoral Camera | Connecting to the monitor via USB or Wifi | Mandatory
69 | HIGIENE | infection control | It should have infection control system with non-retraction valves (Bio system/equivalent). | Mandatory
70 | HIGIENE | EASY HIGIENE instruments and components | DETACHABLE INSTRUMENT TUBING | Mandatory
71 |  |  | REMOVABLE DENTIST’S MODULE HANDLE | Optional
72 |  |  | REMOVABLE LAMP HANDLES | Mandatory
73 |  |  | REMOVABLE LEVERS | Mandatory
74 |  |  | REMOVABLE Spittoon bowl | Mandatory
75 |  |  | REMOVABLE CANNULA GUIDES | Mandatory
76 |  |  | Autosteril, Automatic option for disinfection of the water circuit | Mandatory
77 |  |  | REMOVABLE LEVERS for cleaning and sanitization | Mandatory
78 | POWER SUPPLY | EU plug with ground | 220-240 [V] +/- 10% 50/60 [Hz] | Mandatory
79 | Safety and Standards - Hardware | International Regulations | ISO 9001:2015 Quality management systems Requirements | Mandatory
80 |  | International Regulations | CE certificates or EU declaration of conformity | Mandatory
81 |  | International Regulations | Manufactured conform to the Medical Devices Directive (93/42/EEC), class IIa | Mandatory
82 | Safety and Standards - Hardware | International Regulations | IEC 80601-2-60:2019 Medical electrical equipment - Part 2-60: Particular requirements for the basic safety and essential performance of dental equipment | Mandatory
83 | Warranty and maintenance | Full Service Conditions- no restrictions | Warranty Period: min 2 years | Mandatory
84 | installation and setting of workstations | included for full set | The Stomatologic unit must include all components and accessories for normal operation, Supplier must perform the installation, safety and operation checks before handover. Training of user and techicians must be provided. Equipment assembly and commissioning on site at beneficiary location. Supplier must provide a full operational and fully working working devices | Mandatory
85 | Installation |  | Supplier must perform the installation, safety and operation checks before handover. For this supplier must contact beneficiary before delivery for facility conections coordination of unit. | Mandatory
86 | Installation |  | Equipment assembly and commissioning on site at beneficiary location | Mandatory
87 | Transportation |  | Supplier must include the transportation to the final destination | Mandatory
88 | Documentation |  | User’s manuals, in Romanian language (also in electronic format) | Mandatory
89 | Documentation |  | Service manual, preferable in Romanian language or English (also in electronic form) | Mandatory
90 | Documentation |  | Installation and setup notes / project documentation of delivered state | Mandatory
91 | Local representative/partner | To provide service for the product | Located in Moldova | Mandatory
92 | Local representative/partner | To provide service for the product | Experience in medical equipment procurement and service dental equipment, certificated from manufacturer with presented confirmation | Mandatory
93 | Year of production |  | Must be manufactured not before 2024 year | Mandatory

## D- Pulmonology
Annex D
ITB 2025/EURO/MDA/0002
Procurement of the Interventional Pulmonology workstation
List of requirements and specifications
NAME, CATEGORY AND CODING
1 | Quantity | 1 |  |  |  | Bidder name:
2 | Generic name | Interventional Pulmonology workstation consisting of, electrosurgery, argon-plasma coagulation, cryosurgery CO2. (Statie dedicata bronhoscopiei interventional e formata din Unitatea chirurgicala, Unitatea de coagulare cu Argon-Plasma, Unitatea Criochirurgie cu CO2) |  |  |  | Manufacturer
3 | Level of use | National hospital/ specialized clinics |  |  |  | Model
4 | Place of use | Institute of Oncology |  |  |  | Country of origin:
PURPOSE OF USE
1 | Clinical or other purpose |  |  |  |  | Meet WHO SPECS | NOTE DEVIATION
2 | Overview of functional requirements
Item | Group | Category | Description | Acceptability Criteria | Meet WHO SPECS | Proposed Specificacion (to be completed by the bidder) | Reference from official document / brochure / page
1 | Construction type | Body type and materials
2 | Used for |  | Station dedicated to multiple applications in bronchoscopy and interventional pneumology,especially for minimally invasive endoscopic surgery at the level of the tracheobronchial tree and mediastinum. | Mandatory
3 | Classifications
4 | TECHNICAL CHARACTERISTICS and PHYSICAL CHARACTERISTICS
5 |  | SET OF DEVICES PULMONOLOGY WORKSTATION (manufactured from same manufacturer)
6 |  |  | ELECTROSURGERY UNIT | Mandatory
7 |  |  | ARGON PLASMA COAGULATION (APC) UNIT | Mandatory
8 |  |  | CRYOSURGERY UNIT CO2 | Mandatory
9 |  |  | TROLLEY (STAND) | Mandatory
10 |  |  | CYLINDER CO2
11 |  |  | CYLINDER ARGON | Mandatory
12 |  |  | SET OF ACCSESORIES< COMPONENTS NECESSARY for interconnecting modules and making them work together | Mandatory
13 |  | ELECTROSURGERY UNIT
14 |  | INTENDED FOR | Monopolar applications of tumor recanalization and hemostasis, and more | Mandatory
15 |  | Maximum cutting power | up to 400W at 300 Ohms | Mandatory
16 |  | COAGULATION power | 360W | Mandatory
17 |  | SCREEN | full touch screen or buttons with screen | Mandatory
18 |  | DETECTION FOOT SWITCH | automatic Foot switch recognition once attached | Mandatory
19 |  | CONTROL OF NEUTRAL ELECTRODE | Monitoring the quality of contact with the patient of the neutral electrode | Mandatory
20 |  | OPERATING MODES | cutting and coagulations modes | Mandatory
21 |  |  | Monopolar | Mandatory
22 |  |  | Bipolar | Mandatory
23 |  |  | argon plasma APC | Mandatory
24 |  |  | Endo-Cut endoscopic cutting mode | Mandatory
25 |  |  | Endoscopic APC (argon-plasma coagulation) modes | Mandatory
26 |  | electrode monitoring | electrode monitoring system for safety to prevent burns | Mandatory
27 |  | INCLUDED INS OFFER | Disposable neutral electrode | 200pcs
28 |  |  | Neutral electrode cable minimum 3 meters | 2 pcs
29 |  |  | 2-circuit foot switch for cutting/coagulation control with connection cable included | 2pcs
30 |  |  | Polypectomy loop compatible with use in Endo-Cut mode | 30 pcs
31 |  | ARGON-PLASMA COAGULATION UNIT FOR SURGICAL UNIT
32 |  | SOCKETS | At least one working socket (output channel) of the "plug and play" type with automatic recognition of the instrument. | Mandatory
33 |  | INCLUDED INS OFFER | Probes for endoscopic interventions with the integrated membrane filter (Safety filter) or analog, diameter 1.5 mm, flexible, length 1.5 m with integral filter and instrument recognition, for bronchoscopy, disposable, 10 pcs / set (recommended for use in patients with communicable diseases HIV, TB) | 5 sets
34 |  |  | Adapter for probes | 1 pcs
35 |  | CO2 Cryosurgery Unit
36 |  | INTENDED FOR | Cryosurgery unit for cryobiopsy, cryorecanalization and cryodevitalization in flexible bronchoscopy | Mandatory
37 |  |  | Can be used in endobronchial and transbronchial lung biopsy | Mandatory
38 |  | recognise probes | The unit will recognize the connected instrument and automatically adjust the parameters for each probe type. | Mandatory
39 |  | display on screen | All unit settings visible on the front panel display | Mandatory
40 |  | display on screen | Timecounting function to indicate the activation duration and assist the user with acoustic and visual signals | Mandatory
41 |  | Gas | Cooling gas | CO2
42 |  | Gas | Gas inlet pressure | 45-65bar
43 |  |  | Access to the flexible probe should be ensured and compatible, through a bronchoscope (examples of existing Olympus BF-1TH190, BF-H190, BF-1TQ170, BF-UC190F) | mandatory
44 |  |  | Preset programs | minim 10
45 |  |  | activation pedal (control pedal) | 1
46 |  |  | Flexible cryosurgery probes, disposable, diameter 1.1 mm, length 1150 mm, with protective sheath with length 800±20 mm; for bronchoscope working channel minimum diameter 2.6 mm, (5 pcs / set, sterile packed / per pc) | 5 sets
47 |  |  | Flexible cryosurgery probes, disposable, diameter 1.7 - 1.9 mm, length 1150 mm, (5 pcs / set, sterile packed / per pc) | 5 sets
48 |  |  | Flexible cryosurgery probes, disposable, diameter 2 - 2.4 mm, length 1150 mm, (5 pcs / set, sterile packed / per pc) | 5 sets
49 |  |  | Flexible cryosurgery probes, reusable, diameter 1.7 - 1.9 mm, length 1150 mm, (autoclavable/plasma sterilization/ultrasound). | 2 sets
50 |  | TROLLEY | trolley and installed equipment must be from same manufacturer ( argonand CO2 cylinders can be exception) | Mandatory
51 |  |  | Cilynder holder (for 2 cylinders argon and CO2) | Mandatory
52 |  |  | equiped with basket for accessories | Mandatory
53 |  |  | 4 non static casters, 2 must be with brakes | Mandatory
54 |  |  | included Carrying handle | Mandatory
55 |  | CYLINDER for CO2 | 1xminim 5liters (DIN) or the volume provided by the cylinder should be optimal according to the manufacturer's recommendations | Mandatory
56 |  |  | must be fully adopted with holder from trolley | Mandatory
57 |  |  | CO2 regulator ( DIN) with pressure gauge | Mandatory
58 |  |  | CO2 pipe | Mandatory
59 |  |  | nececary pipes ,conectors and adaptyers | Mandatory
60 |  | CYLINDER for ARGON | 1x minim 5liters (DIN) | Mandatory
61 |  |  | must be fully adopted with holder from trolley | Mandatory
62 |  |  | Argon regulator ( DIN) with pressure gauge | Mandatory
63 |  |  | Pneumatic argon cable (3m and more) | Mandatory
64 |  |  | Argon bacteria filter, sterile ( min 5 pcs) | Mandatory
65 |  |  | nececary pipes, conectors and adapters | Mandatory
66 |  | ACCSESORIES AND COMPONENTS INCLDED | All connection accessories, monopolar cables, neutral electrodes, foot switch | Mandatory
67 | Safety and Standards - Hardware | International Regulations | ISO 9001:2015 Quality management systems — Requirements | Mandatory
68 |  | International Regulations | ISO 9001:2015 Quality management systems — Requirements | Mandatory
69 |  | International Regulations | CE mark/EU declaration of conformity | Mandatory
70 |  | International Regulations | ISO 13485 Medical devices and/or MDR (EU) 2017/745, Class IIB | Mandatory
71 | Safety and Standards - Hardware | INTERNATIONAL REGULATIONS | minimum certificates - IEC 60601-1 E96 General requirements for basic safety and essential performance. - IEC 60601-2-2 Medical Electrical Equipment – Part 2-2: Particular requirements for the safety of high frequency surgical equipment | Mandatory
72 | Warranty and maintenance | Service Conditions- no restrictions | Warranty Period: min 2 years | Mandatory
73 | installation and setting of workstations | included for full set | Supplier must provide a full operational and fully working working devices | Mandatory
74 | Installation | included for full set | The Interventional Pulmonology workstation must include all accessories for normal operation, Supplier must perform the installation, safety and operation checks before handover. Training of user and techicians must be provided. Equipment assembly and commissioning on site at beneficiary location | Mandatory
75 | Transportation | included for full set | Supplier must include the transportation to the final destination | Mandatory
76 | Documentation | included for full set | User’s manuals, in Romanian language (also in electronic format) | Mandatory
77 | Documentation | included for full set | Service manual, preferable in Romanian language or English (also in electronic form) | Mandatory
78 | Documentation | included for full set | Installation and setup notes / project documentation of delivered state | Mandatory
79 | Local representative/partner | To provide service for the product | Located in Moldova | Mandatory
80 | Local representative/partner | To provide service for the product | Experience in medical equipment procurement and service of similar medicals equipment, certificated from manufacturer with presented confirmation | Mandatory
81 | Year of production |  | Must be manufactured not before 2024 year | Mandatory

## E-electrosur unit argplasma
Annex E
ITB 2025/EURO/MDA/0002
Procurement of the Electrosurgical unit with argon-plasma
List of requirements and specifications
NAME, CATEGORY AND CODING
1 | Quantity | 1 |  |  |  | Bidder name:
2 | Generic name | Electrosurgical unit with argon-plasma dedicated for gastrointestinal procedures such as ESD (Electrocoagulator predestinat pentru chirurgia mucoasei tractului digestiv si cailor biliare) |  |  |  | Manufacturer
3 | Level of use | National hospital/ specialized clinics |  |  |  | Model
4 | Place of use | Institute of Oncology |  |  |  | Country of origin:
PURPOSE OF USE
1 | Clinical or other purpose | for use in endoscopic procedures of the digestive tract (TG), biliary tract, duodenum, ERCP |  |  |  | Meet WHO SPECS | NOTE DEVIATION
2 | Overview of functional requirements
Item | Group | Category | Description | Acceptability Criteria | Meet WHO SPECS | Proposed Specificacion (to be completed by the bidder) | Reference from official document / brochure / page
1 | Construction type | Main components
2 |  |  | Electrosurgical unit module | Mandatory
3 |  |  | Argon module unit (wiill be accepted variant if this module can be integrated in electrosurgical module) | Mandatory
4 |  |  | equipment trolley | Mandatory
5 |  |  | electrosurgical instruments | Mandatory
6 |  |  | foot switch | Mandatory
7 |  | HIGH FREQUENCY ELECTROCOAGULATOR UNIT
8 |  | INTENDED FOR | for use in endoscopic procedures of the digestive tract (TG), biliary tract, duodenum | Mandatory
9 |  | WORKING FREQUENCY | High frequency (HF) 440kHz ±25% | Mandatory
10 |  | POWER | Maximum power in HF (high frequency) not more than 120W | Mandatory
11 |  | MONOPOLAR ARGON-PLASMA COAGULATION APC REGIMES, MAPC modes for flexibile probes | minimum 4 | Mandatory
12 |  | SCREEN | full touch screen or buttons with screen | Mandatory
13 |  | DETECTION FOOT SWITCH | automatic Foot switch recognition once attached | Mandatory
14 |  | CONTROL OF NEUTRAL ELECTRODE | Monitoring the quality of contact with the patient of the neutral electrode | Mandatory
15 |  | OPERATING MODES | cutting and coagulations modes | Mandatory
16 |  |  | Monopolar | Mandatory
17 |  |  | Bipolar | Mandatory
18 |  |  | argon plasma coagulation (monopolar APC) | Mandatory
19 |  | AUTOTEST | self-test of generator and accessories after switching the power on, service codes display | Mandatory
20 |  | OVERLOAD | generator overload protection | Mandatory
21 |  | CONSTANT CONTROL | real-time monitoring of all parameters | Mandatory
22 |  | ARGON-PLASMA COAGULATION UNIT FOR SURGICAL UNIT
23 |  | SOCKETS | At least one working socket (output channel) of the "plug and play" type with automatic recognition of the instrument. | Mandatory
24 |  | MODES | Non-contact monopolar argon plasma coagulation for superficial hemostasis, devitalization and ablation with ionized argon gas. | Mandatory
25 |  | MODES | For ESD, duodenal procedures with flexible probes, there should be an automatic mode for detecting the distance to the tissue and the parameters necessary for safe hemostasis, minimal carbonization, limited depth of penetration of the plasma beam for fragile tissues. Supplier must describe the offered mode and possibilities. | Mandatory
26 |  | PROBES | Flexible gastrointestinal probes, argon-plasma coagulation, with automatic recognition, with axial, radial and lateral tip, diameter 2.3mm, length 2.1-2.3 meters | 50 pcs
27 |  | TROLLEY | trolley and installed equipment must be from same manufacturer ( argon cylinders can be exception) | Mandatory
28 |  |  | Cilynder holder (for 1 cylinders minim 1x5l or 1x10l) | Mandatory
29 |  |  | equiped with basket for accessories | Mandatory
30 |  |  | 4 -5 non static casters, 2 must be with brakes | Mandatory
31 |  |  | included Carrying handle | Mandatory
32 |  | CYLINDERS | 1x5liters or 1x10 liter (DIN) | Mandatory
33 |  |  | must be fully adopted with holder from trolley | Mandatory
34 |  |  | Argon regulator ( DIN) with pressure gauge | Mandatory
35 |  |  | Pneumatic argon cable (1m and more) | Mandatory
36 |  |  | Argon bacteria filter, sterile ( min 5 pcs) | Mandatory
37 |  |  | nececary pipes and conectors | Mandatory
38 |  | ACCSESORIES AND COMPONENTS INCLDED | All connection accessories, monopolar cables, neutral electrodes, foot switch | Mandatory
39 | POWER SUPPLY | EU PLUG WITH GROUND | 220-240 [V] +/- 10% 50/60 [Hz] | Mandatory
40 | Software | UPDATES | If any updates or new software during warranty period available must be free installed | Mandatory
41 | Certifications | CE, EU | CE certification or EU declaration of conformity | Mandatory
42 | Safety and Standards - Hardware | INTERNATIONAL REGULATIONS | ISO 9001:2015 Quality management systems — Requirements | Mandatory
43 | Safety and Standards - Hardware | INTERNATIONAL REGULATIONS | ISO 13485 Medical devices and/or MDR (EU) 2017/745, Class IIB | Mandatory
44 | Safety and Standards - Hardware | PROTECTION AGAINST ELECTRIC SHOCK | Class I | Mandatory
45 | Safety and Standards - Hardware | INTERNATIONAL REGULATIONS | minimum certificates - IEC 60601-1 E96 General requirements for basic safety and essential performance. - IEC 60601-2-2 Medical Electrical Equipment – Part 2-2: Particular requirements for the safety of high frequency surgical equipment | Mandatory
46 | Warranty and maintenance | SERVICE CONDITIONS NO RESTRICTIONS | Warranty Period: minim 2 years | Mandatory
47 | installation and setting of workstations | INCLUDED FOR FULL SET | Supplier must provide a full operational and fully working working devices | Mandatory
48 | Installation |  | Supplier must perform the installation, safety and operation checks before handover | Mandatory
49 | Installation |  | Equipment assembly and commissioning on site at beneficiary location | Mandatory
50 | Transportation |  | Supplier must include the transportation to the final destination | Mandatory
51 | Documentation |  | User’s manuals, in Romanian language (also in electronic format) | Mandatory
52 | Documentation |  | Service manual, preferable in Romanian language or English (also in electronic form) | Mandatory
53 | Documentation |  | Installation and setup notes / project documentation of delivered state | Mandatory
54 | Local representative/partner | TO PROVIDE SERVICE FOR THE PRODUCT | Located in Moldova, with maintenance certified engineer from manufacturer | Mandatory
55 | Local representative/partner | TO PROVIDE SERVICE FOR THE PRODUCT | Experience in medical equipment procurement and service, certified from manufacturer | Mandatory
56 | Year of production |  | Must be manufactured not before 2024 year | Mandatory

## F- ultrasound bath
Annex F
ITB 2025/EURO/MDA/0002
Procurement of the Healthcare sonic irrigator (Ultrasound bath)
List of requirements and specifications
NAME, CATEGORY AND CODING
1 | Quantity | 1 |  |  |  | Bidder name:
2 | Generic name | Healthcare sonic irrigator for generalsurgical instruments, Cannulated(lumened)instruments. It is designed to thoroughly clean (remove tissue, blood, and other contaminants) from a variety of reusable surgical instruments. |  |  |  | Manufacturer
3 | Level of use | National hospital/ specialized clinics |  |  |  | Model
4 | Place of use | Institite of Oncology_Sterilization, Disinfection or clinical departments |  |  |  | Country of origin:
PURPOSE OF USE
1 | Clinical or other purpose |  |  |  |  | Meet WHO SPECS | NOTE DEVIATION
2 | Overview of functional requirements
Item | Group | Category | Description | Acceptability Criteria | Meet WHO SPECS | Proposed Specificacion (to be completed by the bidder) | Reference from official document / brochure / page
1 | Construction type | Body type and materials
2 | Used for |  | Healthcare sonic irrigator for generalsurgical instruments, Cannulated(lumened)instruments. It is designed to thoroughly clean (remove tissue, blood, and other contaminants) from a variety of reusable surgical instruments. | Mandatory
3 | Classifications |  | Medicaldevice classI Class 1 electricalshockprotection,externally powered, IP rating:IPX0 Not suitable for use in oxygen-rich environment Designed for continuous use | Mandatory
4 | TECHNICAL CHARACTERISTICS and PHYSICAL CHARACTERISTICS
5 |  | CAPACITY | 25 l-30l | Mandatory
6 |  | FRAME, BODY type | stand alone or tabletop | Optional
7 |  | FRAME, BODY | The equipment vessel is made of AISI 316L stainless steel, the frame is made of non-corrosive materials such as AISI 304/316L or other similar resistant materials. | Mandatory
8 |  | Electronics | minim 4" multi-color screen panel with buttons or touch screen option | Mandatory
9 |  |  | RS232 or USB connection for technical actions (maintenance) | Mandatory
10 |  |  | Ultrasonic system: minimum 5 piezoelectric transducers,operating at 30kHz to 42kHz dependent on load. Minimal 35 Wpower per transducer.
11 |  | Screen | multicolor, LCD, TFT or other | Mandatory
12 |  | Operating | Touch Screen or button | Optional
13 |  | CICLE | Operating process: Ultrasonic activity,irrigation ports,and waterheating. | Mandatory
14 |  |  | Load weigt instruments not less than 3 kg. | Mandatory
15 |  |  | minim 5 pre-set factory programs | Mandatory
16 |  |  | programs that can be programmed by the user | Mandatory
17 |  |  | The cycle progress or number of cicle should be visible on screen | Mandatory
18 |  |  | Type of cleaned instruments: Generalsurgicalinstruments, Cannulated(lumened)instruments | Mandatory
19 |  |  | The instrument basket can be loaded with a mix of general and cannulated instruments on same time. | Mandatory
20 |  |  | The minimum Information on the Display During a Wash Cycle : Currentwashcyclestage, Temperature of fluids , Shows either the time elapsed or the time remaining for the current stage, Progress bar and name ofwash program | Mandatory
21 |  |  | Fast Cycle must be in one of minimal preinstalled programs | Mandatory
22 |  |  | Basket/Tray can hold minimum 8 lumen instruments on same time | Mandatory
23 |  | Irrigation system, Water requirements, temperature monitoring | min.2-port flush manifolds in 2corner of the basket, producing up to 1.5 bar(at a flowrate of 2.5 liters/min. | Mandatory
24 |  |  | Measurement range: +2ºC to +60ºC Temp.display:+5ºC to+60ºC Measurement accuracy : ±0.5ºC. | Mandatory
25 |  |  | Pressure: minimal 2.0 bar to 10-12 bar maximum | Mandatory
26 |  | Detergents and chemicals, dezinfectants used in prcess | integrated place and recipients inside for detergents and chemicals used in cleaning cicles | Mandatory
27 |  |  | disinfectants and chemicals must be provided for minim 50 standard cicles with offered device | Mandatory
28 |  |  | integrated Stainless-steel dosing pumps for chemical products. | Mandatory
29 |  |  | "Open" system for the use of detergents and disinfectants, chemicals. Healthcare sonic irrigator must be open to use different detergents and chemicals, not only from producer of offered MD | Mandatory
30 | Safety | SAFETY and save energy | Class 1 electrical shock protection | Mandatory
31 |  |  | If the ultrasonic cleaner is inactive for more than 10 minutes, the display enters the screensaver mode | Mandatory
32 | ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS | ACCESSORIES | Basket for instruments | Mandatory
33 |  |  | chemical bottles must be present | Mandatory
34 |  | ELECTRICAL CHARACTERISTICS | 220V, 50 Hz, I phase, with earthing, euro style plug | Mandatory
35 | Safety and Standards - Hardware | International Regulations | ISO 9001:2015 Quality management systems — Requirements | Mandatory
36 |  | International Regulations | EN ISO 15883‐1 Washer‐disinfectors | Mandatory
37 |  | International Regulations | Safety Standard IEC 61010 Safety requirements for washing machines and disinfection used to treat medical materials | Mandatory
38 | Warranty and maintenance | Service Conditions -no restrictions | Warranty Period: minim 2 years | Mandatory
39 | installation and setting of workstations | included for full set | Supplier must provide a full operational and fully working working devices | Mandatory
40 | Installation | included for full set | The sonic irrigator must include all accessories for normal operation, Supplier must perform the installation, safety and operation checks before handover. Training of user and techicians must be provided. Equipment assembly and commissioning on site at beneficiary location | Mandatory
41 | Transportation | included for full set | Supplier must include the transportation to the final destination | Mandatory
42 | Documentation | included for full set | User’s manuals, in Romanian language (also in electronic format) | Mandatory
43 | Documentation | included for full set | Service manual, preferable in Romanian language or English (also in electronic form) | Mandatory
44 | Documentation | included for full set | Installation and setup notes / project documentation of delivered state | Mandatory
45 | Local representative/partner | To provide service for the product | Located in Moldova | Mandatory
46 | Local representative/partner | To provide service for the product | Experience in medical equipment procurement and service of cleaning and disinfection equipment, certificated from manufacturer with presented confirmation | Mandatory
47 | Year of production |  | Must be manufactured not before 2024 year | Mandatory

## G- BRAVOS
Annex G
ITB 2025/EURO/MDA/0002
Procurement of the BRAVO
List of requirements and specifications
EXISTING SYSTEM
Set Applicator Brachytherapy cutaneous tumors for BRAVOS
Specification
| Destination | Skin tumor applicator used to extend the possibilities of Varian BRAVOS Brachytherapy apparatus for treating skin cancer by a less invasive, quicker methods that offer better cosmetic results for pacients
Item | The components: |  |  | Acceptability Criteria | Meet WHO SPECS | Proposed Specificacion (to be completed by the bidder) | Reference from official document / brochure / page
1 | Surface Applicator Set (GM11004590) |  | 1 unit
2 | Catheter Flap Applicator set (GM11004370) |  | 2 unit | Mandatory
3 | On site expert training |  | 1 day | Mandatory
4 | Team (radiation oncologist+ medical physicist) clinical training in a Hospital where are using the same treatment technique |  | 5 days | Mandatory
5 | Surface Applicator Set must be manufactured from same manufacturer as existing BRAVOS |  |  | Mandatory
6 | International Regulations |  | standards and regulations similirar as existing unit | Mandatory
7 | Service Conditions no restrictions |  | Warranty Period: minim 2 years | Mandatory
8 | included for full set |  | Supplier must provide a full operational and fully working working devices adjusted to existing unit | Mandatory
9 | included for full set |  | must include all accessories for normal operation, Supplier must perform the installation, safety and operation checks before handover. Training of user and techicians must be provided acordin point 4 from list. Equipment assembly and commissioning on site at beneficiary location | Mandatory
10 | included for full set |  | Supplier must include the transportation to the final destination | Mandatory
11 | included for full set |  | User’s manuals, in Romanian language (also in electronic format) | Mandatory
12 | included for full set |  | Service manual, preferable in Romanian language or English (also in electronic form) | Mandatory
13 | included for full set |  | Installation and setup notes / project documentation of delivered state | Mandatory
14 | To provide service for the product |  | Located in Moldova | Mandatory
15 | To provide service for the product |  | Experience in medical equipment procurement and service of equipment, certificated from manufacturer with presented confirmation | Mandatory
16 | To provide service for the product |  | Must be manufactured not before 2024 year | Mandatory

## H- ultrasonic surgery
Annex H
ITB 2025/EURO/MDA/0002
Procurement of the Ultrasonic Scalpel
List of requirements and specifications
NAME, CATEGORY AND CODING
1 | Quantity | 1 |  |  |  | Bidder name:
2 | Generic name | Ultrasonic Scalpel Ultrasound-based surgical system with Harmonic and bipolar ultrasound cutting (dissection) and coagulation functions |  |  |  | Manufacturer
3 | Level of use | National hospital/ specialized clinics |  |  |  | Model
4 | Place of use | Institute of Oncology |  |  |  | Country of origin:
PURPOSE OF USE
1 | Clinical or other purpose | Thoracic surgery: radical surgery for esophageal cancer, radical surgery for lung cancer, mediastinal surgery |  |  |  | Meet WHO SPECS | NOTE DEVIATION
2 | Overview of functional requirements | Electrosurgical device with instruments/handpieces intended for tissue dissection and coagulation.
Item | Group | Category | Description | Acceptability Criteria | Meet WHO SPECS | Proposed Specificacion (to be completed by the bidder) | Reference from official document / brochure / page
1 | Construction type | Body type and materials
2 | Used for |  | Electrosurgical device with instruments/handpieces intended for tissue dissection and coagulation. | Mandatory
3 | Classifications | Main unit | Control console (generator) | Mandatory
5 |  | Interface | screen | Mandatory
6 |  | parameters | Operating principle: based on "Enseal" and "Harmonic" technologies | Mandatory
7 |  | parameters | Ultrasonic dissection and coagulation function | Mandatory
8 |  | Working frequency | 55KHz ± 2KHz | Mandatory
9 |  | coagulations modes | must include Bipolar coagulation function | Mandatory
10 |  | Advanced output | Advanced bipolar output without the need for a neutral electrode | Mandatory
11 |  | dissection and coagulation of vessels | Possibility to perform dissection and coagulation of vessels with diameter/thickness | Mandatory
12 |  | Ability to adjust the amplitude of the nozzle/accessory oscillations for ultrasonic dissection/coagulation | minimum 5 levels | Mandatory
13 |  | tissue adaptive technology | The presence of tissue adaptive technology that will enable greater surgical precision and performance by dynamically optimizing energy delivery in response to changing physical characteristics of the tissue | Mandatory
14 |  | recognition of connected instruments | Automatic recognition of connected instruments to minimize setup time | Mandatory
15 |  | self-test system | Generator self-test system, ability to diagnose correct instrument connection with notification system and alarms with visual and audible indicator | Mandatory
16 |  | Screen | color screen | Mandatory
17 |  | instrument control | Possibility of activating the instruments manually and with a foot switch | Mandatory
18 |  | type surgeries | For open and endoscopic surgeries | Mandatory
19 |  | Accessory:
20 |  | Stand (trolley) | Dedicated wheeled support (stand) for installation and securing the generator, minimum 4 wheels, minimum 1 with brake. | 1 unit
21 |  | ACCSESORIES AND COMPONENTS | Foot switch with cable | 1
22 |  | ACCSESORIES AND COMPONENTS | Handpiece compatible with Multifunctional Instruments with Ultrasonic Coagulation and Dissection for Endoscopic and Open Surgery (e.g. type HP054) | 3
23 |  | ACCSESORIES AND COMPONENTS | Handpiece compatible with Multifunctional Instruments with Ultrasonic Coagulation and Dissection for Open Surgery (e.g. HPBLUE type) | 1
24 |  | ACCSESORIES AND COMPONENTS | Adapter for Harmonic ultrasonic instruments | 1
25 |  | ACCSESORIES AND COMPONENTS | Multifunctional instrument for ultrasonic coagulation and dissection, used in endoscopic and open surgery, diameter 5 mm, length 20-25 cm, seal-dissection blade length 14-15 mm, curved jaws. Possibility of 360° rotation. Presence of tissue-adaptive technology, coagulation of vessels up to 7 mm in diameter, and at least 3 buttons with energy level (power) control functions | 12
26 |  | ACCSESORIES AND COMPONENTS | Multifunctional instrument for ultrasonic coagulation and dissection, used in endoscopic and open surgery, diameter 5 mm, length 32-37 cm, seal-dissection blade length 14-15 mm, curved jaws. Possibility of 360° rotation. Presence of tissue-adaptive technology, coagulation of vessels up to 7 mm in diameter, and at least 3 buttons with energy level (power) control functions | 6
27 |  | ACCSESORIES AND COMPONENTS | Multifunctional instrument for ultrasonic coagulation and dissection, scissors type, used in open surgery, length 16-20 cm, seal-dissection blade length 15-16 mm, curved jaws. Coagulation of vessels up to 5 mm in diameter, and at least 2 buttons with energy level (power) control functions | 6
28 |  | ACCSESORIES AND COMPONENTS | Multifunctional instrument for cutting and sealing bipolar blood vessels, used in endoscopic surgery, with the possibility of articulation up to 110 degrees, diameter 5 mm, length 32-37 cm, curved jaws, with minimum 2.5 meter interconnecting cable. Coagulation of vessels up to 7 mm in diameter. | 6
29 |  | ACCSESORIES AND COMPONENTS | Multifunctional instrument for cutting and sealing bipolar blood vessels, used in open surgery, diameter 10-13 mm, length 18-23 cm, wide curved jaws, with minimum 2.5 meter interconnecting cable. Coagulation of vessels up to 7 mm in diameter. | 6
30 | POWER SUPPLY | EU plug with ground | 220-240 [V] +/- 10% 50/60 [Hz] | Mandatory
31 | Safety and Standards | International Regulations | ISO 9001:2015 Quality management systems — Requirements | Mandatory
32 |  | International Regulations | ISO 13485 | Mandatory
33 |  | International Regulations | CE marking or EU declaration of conformity | Mandatory
34 | Safety and Standards - Hardware | International Regulations | minimum certificates - IEC 60601-1 E96 General requirements for basic safety and essential performance. - IEC 60601-2-2 Medical Electrical Equipment – Part 2-2: Particular requirements for the safety of high frequency surgical equipment | Mandatory
35 | Warranty and maintenance | Service Conditions -no restrictions | Warranty Period: minim 2 years | Mandatory
36 | installation and setting of workstations | included for full set | Supplier must provide a full operational and fully working working devices | Mandatory
37 | Installation | included for full set | The Ultrasonic surgery unit must include all accessories for normal operation, Supplier must perform the installation, safety and operation checks before handover. Training of user and techicians must be provided. Equipment assembly and commissioning on site at beneficiary location | Mandatory
38 | Transportation | included for full set | Supplier must include the transportation to the final destination | Mandatory
39 | Documentation | included for full set | User’s manuals, in Romanian language (also in electronic format) | Mandatory
40 | Documentation | included for full set | Service manual, preferable in Romanian language or English (also in electronic form) | Mandatory
41 | Documentation | included for full set | Installation and setup notes / project documentation of delivered state | Mandatory
42 | Local representative/partner | To provide service for the product | Located in Moldova | Mandatory
43 | Local representative/partner | To provide service for the product | Experience in medical equipment procurement and service of medical equipment, certificated from manufacturer with presented confirmation | Mandatory
44 | Year of production |  | Must be manufactured not before 2024 year | Mandatory

## 16-Dispo.padeles ad.ped. (hidden)
Item Disposable padeles for adults and pediatricc