## Lot 1
Annex 2
Technical compliance table of the medical equipment for the Medical Service of the Ministry of Internal Affairs of the Republic of Moldova
|  |  | REQUIRED |  |  | OFFERED
Item |  |  |  |  | Quantity | Item |  | Quantity
LOT 1 | ANESTHESIA APPARATUS Type 1 |  | Composition, accessories, spare parts, consumables |  | 1 | Composition, accessories, spare parts, consumables |  | 1
1.1. | Composition | 1 | Main Unit |  | 1 unit
|  | 2 | Vaporizer, Sevoflurane | Vaporizer, Sevoflurane | 1 unit
|  | 3 | Anesthesia ventilator | Anesthesia ventilator | 1 unit
|  | 4 | Anesthesia monitor | Anesthesia monitor | 1 unit
1.2. | Accessories | 1 | Pressure tube, for O2 and N2O |  | 1 set
|  | 2 | Corrugated tube for surplus gas removal connect. to AGS |  | 1 pc.
|  | 3 | Breathing circuit, reusable, adult |  | 10 pc.
|  | 4 | Y-adapter, plastic |  | 1 set
|  | 5 | Anesthesia mask, S, M, and L |  | 1 set
|  | 6 | Test lung(1.5 L or more) |  | 1 set
|  | 7 | Breathing bag, 3L or more |  | 1 set
|  | 8 | Laryngoscope set |  | 1 set
|  | 9 | Charging station for laryngoscopes |  | 1 pc.
|  | 10 | Self-inflating bag for manual ventilation, for adult (1500 mL), reusable |  | 1 pc.
|  | 11 | Reusable silicone mask size 3 |  | 1 pc.
|  | 12 | Reusable silicone mask size 4 |  | 1 pc.
|  | 13 | Reusable silicone mask size 5 |  | 1 pc.
|  | 14 | Head strap, reusable, silicone, small |  | 1 pc.
|  | 15 | Head strap, reusable, silicone, large |  | 1 pc.
|  | 16 | Guedel airway, size 5 |  | 1 pc.
|  | 17 | Guedel airway, size 4 |  | 1 pc.
|  | 18 | Guedel airway, size 3 |  | 1 pc.
|  | 19 | Guedel airway, size 2 |  | 1 pc.
|  | 20 | Guedel airway, size 1.5 |  | 1 pc.
|  | 21 | Guedel airway, size 0 |  | 1 pc.
|  | 22 | Guedel airway, size 00 |  | 1 pc.
|  | 23 | Reusable CO2 absorber |  | 1 pc.
|  | 24 | Standard configuration 115V/230V |  | 1 pc.
|  | 25 | Patient monitor mounting kit |  | 1 set
|  | * | Any other accessories which is required for installation and operation should be included and described on the offered specification.
2. | Technical Specifications | 2.1 | Main Unit | specifications |  | Main Unit | specifications
|  | a) | Gas mixing control | Electronic
|  | b) | Flow meter | O2 0.3 to 40.0 L/min or wider
N2O 0.5 to 20.0 L/min or wider
|  | c) | Vaporizer hook | 2pcs. or more
|  | d) | Canister for CO2 absorber | Equipped
|  | e) | Oxygen monitor | -
|  | f) | Type | Paramagnetic sensor or fuel cell
|  | g) | Measuring range | 0 to 99% or wider
|  | 2.2 | Vaporizer | specifications |  | Vaporizer | specifications
|  | a) | Kind of agent | Isoflurane, Sevoflurane, Desflurane
|  | b) | Type | Mechanical or Electronical control | Mechanical or Electronical control type
|  | c) | Capacity | 300mL or more | 300mL or more
|  | 2.3 | Anesthesia ventilator | specifications |  | Anesthesia ventilator | specifications
|  | a) | Type | Electronically controlled
|  | b) | Operating mode | Volume and Pressure control, SIMV, VCV, PCV, PSV or CPAP or more
|  | c) | Breathing frequency | 4 to 90 bpm or wider
|  | d) | Tidal volume | 20 to 1900 mL or wider | 20 to 1900 mL or wider
|  | e) | I:E ratio | 4:1 to 1:10 or wider | 4:1 to 1:10 or wider
|  | f) | Peep | 0 to 30 cm H2O or wider | 0 to 30 cmH2O or wider
|  | 2.4 | Anesthesia monitor | specifications |  | Anesthesia monitor | specifications
|  | a) | Parameters | Inspiratory concentration of O2, CO2, N2O, MAC
Breathing frequency, Tidal volume
Minutes volume, Peak pressure, etc.
|  | b) | Size | 10 inch or more, touch screen | 10 inch or more, touch screen
|  | c) | Connection | Input port for patient monitor | Input port for patient monitor
|  | d) | Alarm signal output | Various signals are transferred to Patient monitor
|  | e) | Interface | RS-232C | Various signals are transferred to Patient
|  | 2.5 | Laryngoscope | specifications |  | Laryngoscope | specifications
|  |  | Handle | Rechargeable | Rechargeable
|  |  | Blade type | Macintosh | Macintosh
|  |  | Blade size | 5 sizes | 5 sizes
|  | 2.6 | Power requirement | specifications |  | Power requirement | specifications
|  | a) | AC 220V±10%, single phase, 50Hz
3. | Spare parts | 1 | N/a
4. | Consumables | 1 | Connector for catheter |  | 10 pcs.
|  | 2 | Breathing filter set, Adult |  | 10 sets
|  | 3 | Water trap for Line CO2 probe |  | 12 pcs.
|  | 4 | Line for CO2 probe (10 pcs./box) |  | 12 boxes
|  | 5 | Soda lime sorbent for CO2 |  | 2 pcs.
|  | 6 | Flow sensor (5 pcs./box) |  | 1 box
|  | 7 | Spare bulb for laryngoscope blade |  | 2 sets
* Please add composition, accessories, spare parts, consumables (if applicable)

## Lot 2
Annex 2
Technical compliance table of the medical equipment for the Medical Service of the Ministry of Internal Affairs of the Republic of Moldova
|  |  | REQUIRED |  |  | OFFERED
Item |  |  |  |  | Quantity | Item |  | Quantity
LOT 2 | ANESTHESIA APPARATUS Type 2 |  | Composition, accessories, spare parts, consumables |  | 1 | Composition, accessories, spare parts, consumables |  | 1
1.1. | Composition | 1 | Main Unit |  | 1 unit
|  | 2 | Vaporizer, Sevoflurane | Vaporizer, Sevoflurane | 1 unit
|  | 3 | Anesthesia ventilator | Anesthesia ventilator | 1 unit
|  | 4 | Anesthesia gas monitor | Anesthesia monitor | 1 unit
1.2. | Accessories | 1 | Pressure tube, for O2 and N2O |  | 1 set
|  | 2 | Corrugated tube for surplus gas removal connect. to AGS |  | 1 pc.
|  | 3 | Breathing circuit, reusable, adult |  | 10 pc.
|  | 4 | Y-adapter, plastic |  | 1 set
|  | 5 | Anesthesia mask, S, M, and L |  | 1 set
|  | 6 | Test lung(1.5 L or more) |  | 1 set
|  | 7 | Breathing bag, 3L or more |  | 1 set
|  | 8 | Laryngoscope set |  | 1 set
|  | 9 | Charging station for laryngoscopes |  | 1 pc.
|  | 10 | Self-inflating bag for manual ventilation, for adult (1500 mL), reusable |  | 1 pc.
|  | 11 | Reusable silicone mask size 3 |  | 1 pc.
|  | 12 | Reusable silicone mask size 4 |  | 1 pc.
|  | 13 | Reusable silicone mask size 5 |  | 1 pc.
|  | 14 | Head strap, reusable, silicone, small |  | 1 pc.
|  | 15 | Head strap, reusable, silicone, large |  | 1 pc.
|  | 16 | Guedel airway, size 5 |  | 1 pc.
|  | 17 | Guedel airway, size 4 |  | 1 pc.
|  | 18 | Guedel airway, size 3 |  | 1 pc.
|  | 19 | Guedel airway, size 2 |  | 1 pc.
|  | 20 | Guedel airway, size 1.5 |  | 1 pc.
|  | 21 | Guedel airway, size 0 |  | 1 pc.
|  | 22 | Guedel airway, size 00 |  | 1 pc.
|  | 23 | Reusable CO2 absorber |  | 1 pc.
|  | 24 | Patient monitor mounting kit |  | 1 set
|  | * | Any other accessories which is required for installation and operation should be included and described on the offered specification.
2. | Technical Specifications | 2.1. | Main Unit | REQUIRED specifications |  | Main Unit | OFFERED specifications
|  | a) | Gas mixing control | Analog or Electronic
|  | b) | Flow meter | O2 0.2 to 10.0 L/min or wider | O2 0.3 to 40.0 L/min or wider
|  |  |  | N2O 0.5 to 10.0 L/min or wider | N2O 0.5 to 20.0 L/min or wider
|  | c) | Vaporizer hook | 2pcs. or more | 2pcs. or more
|  | d) | Canister for CO2 absorber | Equipped | Equipped
|  | e) | Oxygen monitor | -
|  | f) | Type | Galvanic sensor or fuel cell | Paramagnetic sensor or fuel cell
|  | g) | Measuring range | 0 to 99% or wider | 0 to 99% or wider
|  | 2.2 | Vaporizer | REQUIRED specifications |  | Vaporizer | OFFERED specifications
|  | a) | Kind of agent | Isoflurane, Sevoflurane, Desflurane
|  | b) | Type | Mechanical or Electronical control type
|  | c) | Capacity | 300mL or more | 300mL or more
|  | 2.3 | Anesthesia ventilator | REQUIRED specifications |  | Anesthesia ventilator | OFFERED specifications
|  | a) | Type | Electronically controlled type | Electronically controlled type
|  | b) | Operating mode | Volume and Pressure control, PS, SIMV or more
|  | c) | Breathing frequency | 4 to 60 bpm or wider
|  | d) | Tidal volume | 20 to 1400 mL or wider | 20 to 1400 mL or wider
|  | e) | I:E ratio | 2:1 to 1:8 or wider | 2:1 to 1:8 or wider
|  | f) | Peep | 0 to 20 cm H2O or wider | 0 to 20 cmH2O or wider
|  | 2.4 | Anesthesia monitor | REQUIRED specifications |  | Anesthesia monitor
|  | a) | Parameters | Inspiratory concentration of O2, CO2, N2O
Breathing frequency, Tidal volume
Minutes volume, Peak pressure, etc.
|  | b) | Size | 10 inch or more, touch screen | 10 inch or more, touch screen
|  | c) | Connection | Input port for patient monitor | Input port for patient monitor
|  | d) | Alarm signal output | Various signals are transferred to Patient monitor
|  | e) | Interface | RS-232C | Various signals are transferred to Patient
|  | 2.5 | Laryngoscope | REQUIRED specifications |  | Laryngoscope | OFFERED specifications
|  |  | Handle | Rechargeable | Rechargeable
|  |  | Blade type | Macintosh | Macintosh
|  |  | Blade size | 5 sizes | 5 sizes
|  | 2.6 | Power requirement | REQUIRED specifications |  | Power requirement | OFFERED specifications
|  | a) | AC 220V±10%, single phase, 50Hz
3 | Spare parts | 1 | N/a
1.3. | Consumables | 1 | Connector for catheter |  | 10 pcs.
|  | 2 | Breathing filter set, Adult |  | 10 sets
|  | 3 | Water trap for Line CO2 probe |  | 12 pcs.
|  | 4 | Line for CO2 probe (10 pcs./box) |  | 12 boxes
|  | 5 | Soda lime sorbent for CO2 |  | 2 pcs.
|  | 6 | Flow sensor (5 pcs./box) |  | 1 box
|  | 7 | Spare bulb for laryngoscope blade |  | 2 sets
* Please add composition, accessories, spare parts, consumables (if applicable)

## Lot 3
Annex 2
Technical compliance table of the medical equipment for the Medical Service of the Ministry of Internal Affairs of the Republic of Moldova
|  |  | REQUIRED |  |  | OFFERED
|  |  | Item |  | Quantity | Item |  | Quantity
LOT 3 | VENTILATOR FOR ADULT |  | Composition, accessories, spare parts, consumables |  | 1 | Composition, accessories, spare parts, consumables |  | 1
1.1. | Composition | 1 | Main unit |  | 1 unit
|  | 2 | Stand |  | 1 unit
|  | 3 | External Air compressor (if necessary) |  | 1 unit
1.2. | Accessories | 1 | Oxygen high pressure hose |  | 1 pc.
|  | 2 | Compressed air high pressure hose (if necessary) |  | 1 pc.
|  | 3 | Respiratory circuit for adult, reusable |  | 3 sets
|  | 4 | Circuit hanger |  | 1 set
|  | 5 | Humidifier with heater |  | 1 set
|  | 6 | Test lung for adult |  | 1 pc.
|  | * | Any other accessories which is required for installation and operation should be included and described on the offered specification.
2. | Technical Specifications | 2.1 | Main Unit | specifications |  | Main Unit | specifications
|  | a) | Applicable patient | Adult or body weight over 5 kg
|  | b) | Control principle | Microprocessor control, based volume and pressure limited ventilation
|  | c) | Ventilation mode | SIMV, PS and CPAP or more
|  | d) | PEEP/CPAP | 1 to 35 cmH2O or wider
|  | e) | Pressure Support | 2 to 35 cmH2O or wider
|  | f) | Respiratory rate | 4 to 80 bpm or wider
|  | g) | Tidal volume | 150 to 2000 mL or wider
|  | h) | Inspiratory time | 0.2 to 3.0 second or wider
|  | i) | I/E ratio | 3:1 to 1:10 or wider
|  | j) | Oxygen concentration | 21 to 100%
|  | k) | Alarm | High and low pressure, Gas supply source failure, Electrical power failure, apnea or more
|  | l) | Indication | Tidal volume, inspiration time, I/E ratio or more
|  | m) | Real-time curves | Pressure, Flow, P/V curve or more
|  | n) | Trigger Sensitivity | 1 to 10 L/min or wider
|  | o) | External gas requirement | Oxygen and Compressed Air (If required)
|  | p) | Power requirement | AC 220V±10%, single phase, 50Hz
|  | 2.2 | Stand | specifications |  | Stand | specifications
|  | a) | Equipped with casters
3. | Spare parts | 1 | N/a
4 | Consumables | 1 | Paper for humidifier (100 pcs./pack) |  | 1 pack
* Please add composition, accessories, spare parts, consumables (if applicable)

## Lot 4
Annex 2
Technical compliance table of the medical equipment for the Medical Service of the Ministry of Internal Affairs of the Republic of Moldova
|  |  | REQUIRED |  |  | OFFERED
Item |  |  |  |  | Quantity | Item |  | Quantity
LOT 4 | INTUBATION FIBROLARYNGOSCOPE |  | Composition, accessories, spare parts, consumables |  | 1 | Composition, accessories, spare parts, consumables |  | 1
1.1. | Composition | 1 | Flexible bronchoscope for laryngoscopy |  | 1 unit
1.2. | Accessories | 1 | Instruments tray | Instruments tray | 1 pc.
|  | 2 | Cleaning brush | Cleaning brush | 1 pc.
|  | 3 | Cleaning brush for valves | Cleaning brush for valves | 1 pc.
|  | 4 | Cleaning kit | Cleaning kit | 1 pc.
|  | 5 | Light source | Light source | 1 pc.
|  | * | Any other accessories which is required for installation and operation should be included and described on the offered specification.
2. | Technical Specifications | 2.1 | Flexible Bronchoscope | specifications |  | Main unit | specifications
|  | a) | Applicable patient | Adult
|  | b) | Viewing direction | 0° | 0°
|  | c) | Field of view | 110° or more | 110° or more
|  | d) | Observation range | 3 to 50 mm or wider | 3 to 50 mm or wider
|  | e) | Distal end diameter | 5.5 mm or less | 5.5 mm or less
|  | f) | Bending capability | Up 180°/ Down 90° or more | Up 180°/ Down 90° or more
|  | g) | Working length | 600 mm or more | 600 mm or more
|  | h) | Forceps channel | 2.0 mm to 2,5 mm | 2.0 mm to 2,5 mm
|  | i) | Air and water nozzle | 1 pc. | 1 pc.
3. | Spare parts | 1 | N/a
4. | Consumables | 1 | Lens cleaner | Lens cleaner | 1 pc.
|  | 2 | Silicon oil | Silicon oil | 1 pc.
|  | 3 | Lithium-ion battery | Silicon oil | 6 pcs.
* Please add composition, accessories, spare parts, consumables (if applicable)
* Please add composition, accessories, spare parts, consumables (if applicable)

## Lot 5
Annex 2
Technical compliance table of the medical equipment for the Medical Service of the Ministry of Internal Affairs of the Republic of Moldova
|  |  |  | REQUIRED |  |  | OFFERED
Item | [5 empty] | Quantity | Item |  | Quantity
LOT 4 | LOT 5 | PATIENT MONITOR |  | Composition, accessories, spare parts, consumables |  | 1 | Composition, accessories, spare parts, consumables |  | 1
1.1. | 1.1. | Composition | 1 | Main Unit |  | 1 unit
|  |  | 2 | Recorder module |  | 1 unit
1.2. | 1.2. | Accessories | 1 | Rechargeable battery pack |  | 1 set
|  |  | 2 | ECG connection cord |  | 1 set
|  |  | 3 | Electrode lead, for 3 electrodes |  | 1 set
|  |  | 4 | Temperature probe, catheter or rectal type, adult |  | 1 set
|  |  | 5 | Temperature probe, disk type, 10 mm dia |  | 1 set
|  |  | 6 | SpO2 connection cord |  | 1 set
|  |  | 7 | SpO2 finger probe |  | 1 set
|  |  | 8 | SpO2 multi-site probe |  | 1 set
|  |  | 9 | NIBP air hose for adult/child, 3.5m |  | 1 set
|  |  | 10 | NIBP cuff for adult |  | 1 set
|  |  | * | Any other accessories which is required for installation and operation should be included and described on the offered specification.
2. | 2. | Technical Specifications | 2.1 | Main unit | specifications |  | Main Unit | specifications
|  |  | a) | Measuring parameters | Arterial Pressure, Pulse, ECG, Respiration, Temperature, SpO2, Non-invasive blood pressure, CO2, Invasive blood continuous cardiac output or Estimated integrated 24 trending CO (tabular/graphical) |  | Measuring parameters
|  |  | b) | Display |  |  | Display
|  |  |  | - Display | Color 10 inch or more
|  |  |  | - Number of waveforms | 15 or more
2. |  |  |  | - Sweep speed | 6.25, 12.5, 25 or 50 mm/s and more
|  |  |  | - Freeze function | Equipped
|  |  | c) | ECG |  |  | ECG
|  |  |  | - ECG lead | 3-electrodes, I, II, III and more
|  |  |  | - Heart rate counting range | 0, 15 to 250 bpm or wider
|  |  |  | - Respiration Measuring range | 0,4 to 125 rpm or wider
|  |  | d) | Non-invasive blood pressure |  |  | Non-invasive blood pressure
|  |  |  | - Measuring method | Oscillometric method
|  |  |  | - Cuff pressure display range | 30 to 280 mmHg or wider
|  |  | e) | Temperature measuring range | 0,2 to 45°C or wider |  | Temperature measuring range
|  |  | f) | SpO2 measuring range | 1 to 100% or wider |  | SpO2 measuring range
|  |  | g) | IBP measuring | -50 to 300 mmHg or wider |  | IBP measuring
|  |  | h) | Carbon Dioxide |  |  | Carbon Dioxide
|  |  |  | - Calculation method | Semi-quantitative or Sidestream
|  |  |  | - Measuring range | 0 to 100 mmHg or wider
|  |  | i) | Recorder | Thermal type |  | Recorder
|  |  |  | - Number of channels | 3 or more
|  |  | j) | Power requirement | AC 220V±10%, single phase, 50Hz |  | Power
| 3. | Spare parts | 1 | ECG connection cord |  | 1 set
|  |  | 2 | Electrode lead, for 3 electrodes |  | 1 set
|  |  | 3 | Temperature probe, catheter type, adult |  | 1 set
|  |  | 4 | Temperature probe, catheter type, child |  | 1 set
|  |  | 5 | Temperature probe, disk type, 10 mm dia |  | 1 set
|  |  | 6 | SpO2 finger probe |  | 1 set
|  |  | 7 | SpO2 multi-site probe |  | 1 set
|  |  | 8 | SpO2 connection cord |  | 1 set
|  |  | 9 | NIBP air hose for adult/child, 3.5m |  | 1 set
|  |  | 10 | NIBP cuff for adult |  | 1 set
| 4. | Consumables | 1 | Disposable electrode, adult, 150 pcs./set |  | 6 sets
|  |  | 2 | Temperature insulation pad, 60 pcs./set |  | 3 sets
|  |  | 3 | Temperature probe cover, 10 pcs./set |  | 3 sets
|  |  | 4 | Recording paper, 10 pcs./set, 50mm x 20m |  | 3 sets
* Please add composition, accessories, spare parts, consumables (if applicable)

## Lot 6
Annex 2
Technical compliance table of the medical equipment for the Medical Service of the Ministry of Internal Affairs of the Republic of Moldova
|  |  | REQUIRED |  |  | OFFERED
Item |  |  |  |  | Quantity | Item |  | Quantity
LOT 6 | OPERATION TABLE |  | Composition, accessories, spare parts, consumables |  | 1 | Composition, accessories, spare parts, consumables |  | 1
1.1. | Composition | 1 | Main unit |  | 1 unit
|  | 2 | Hand remote control unit | Hand remote control unit | 1 pc.
1.2. | Accessories | 1 | Arm board with mattress |  | 2
|  | 2 | Body support | Body support | 2
|  | 3 | Extra radiographic table top plate | Extra radiographic table top plate | 1
|  | 4 | Head rest with attachment | Head rest with attachment | 1
|  | 5 | Knee rest, extension board | Knee rest, extension board | 1
|  | 6 | Levitator | Levitator | 1
|  | 7 | Locking type arm board with pad | Locking type arm board with pad | 1
|  | 8 | L-shaped screen frame with clamp | L-shaped screen frame with clamp | 1
|  | 9 | Screen frame, adjustable type with clamp | Screen frame, adjustable type with clamp | 1
|  | 10 | Urology tray with hose and receptacle | Urology tray with hose and receptacle | 1
|  | 11 | V-attachment | V-attachment | 1
|  | 12 | Gynecological support | Gynecological support | 1
|  | 13 | Perfusion support | Perfusion support | 1
|  | * | Any other accessories which is required for installation and operation should be included and described on the offered specification.
2. | Technical Specifications | 2.1 | Main unit | specifications |  | Main Unit | specifications
|  | a) | Table top size | 2,030mm x 500mm or more
|  | b) | Table top material | Radiolucent plastic plate and/or metal frame
|  |  | Radiolucent area | 1600mm or more
|  | c) | Table top positioning
|  |  | Cummulative Sliding | 310 mm or more
|  |  | Height adjustment (with padding) | 620mm - 1100mm or wider
|  |  | Trendelenburg / Reverse | 30degree / 30degree or more
|  |  | Lateral tilt right / left | 20degree / 20degree or more
|  |  | Back section up / down | 60degree / 40degree or more, detachable
|  |  | Head section up / down | 45degree / 90degree or more, manual, detachable
|  |  | Leg section up / down | 15degree / 90 degree or more, detachable, pair
|  | d) | X-ray cassette tray size | 390 x 288mm or bigger
|  | e) | Maximum load | 450kg for elevation, 250kg for all positions or more
|  | f) | Return to level | One button return to zero level or per each position separately
Flex/Reflex or more
|  | g) | Controls | Hand Remote control unit with or On table colums and wired hand control unit secondary switch
|  | h) | Mobility | 4 swivel casters and electro-hydraulic
4-point base
|  | j) | Built-in battery and charger | DC 24V
|  |  | Power requirement | AC 220V±10%, Single phase, 50Hz
|  | 2.2 | Mayo instrument tray with stand | specifications |  | Main Unit | specifications
|  | a) | Tray size | 350(W) x 500(D) mm or more
|  | b) | Lift range | 850-1200 mm or wider
|  | c) | Material | Stainless steel
3. | Spare parts | 1 | N/a
4. | Consumables | 1 | N/a
* Please add composition, accessories, spare parts, consumables (if applicable)

## Lot 7
Annex 2
Technical compliance table of the medical equipment for the Medical Service of the Ministry of Internal Affairs of the Republic of Moldova
|  |  | REQUIRED |  |  | OFFERED
Item |  |  |  |  | Quantity | Item |  | Quantity
LOT 7 | SUCTION UNIT |  | Composition, accessories, consumables, spare parts |  | 2 | Composition, accessories, consumables, spare parts |  | 2
1.1. | Composition | 1 | Main unit |  | 2 unit
1.2. | Accessories | 1 | Suction bottle | Suction bottle | 4 pcs.
|  | 2 | Suction tube | Suction tube | 4 pcs.
|  | * | Any other accessories which is required for installation and operation should be included and described on the offered specification.
2. | Technical Specifications | 2.1 | Main Unit | specifications |  | Main unit | specifications
|  | a) | Type | 1 bottle type or more
|  | b) | Maximum Vacuum | 650 mmHg or more | 650 mmHg or more
|  | c) | Flow rate | 30 L / min or more | 30 L / min or more
|  | d) | Capacity of suction bottle | Total 4 L or more, autoclavable | Total 4 L or more, autoclavable
|  | e) | Caster | Equipped, with brakes | Equipped, with brakes
|  | f) | Power requirement | AC 220V±10%, single phase, 50Hz | AC 220V±10%, single phase, 50Hz
1.4. | Spare parts | 1 | Suction bottle | Suction bottle | 4 pcs.
|  | 2 | Suction tube | Suction tube | 2 pcs.
1.3. | Consumables | 1 | Bacteria Filter |  | 20 pcs.
* Please add composition, accessories, spare parts, consumables (if applicable)
Note on F6: LL: Quantity is set for 2 units of equipment

## Lot 8
Annex 2
Technical compliance table of the medical equipment for the Medical Service of the Ministry of Internal Affairs of the Republic of Moldova
|  |  | REQUIRED |  |  | OFFERED
Item |  |  |  |  | Quantity | Item |  | Quantity
LOT 8 | ELECTROSURGICAL UNIT |  | Composition, accessories, consumables, spare parts |  | 2 | Composition, accessories, consumables, spare parts |  | 2
1.1. | Composition | 1 | Main unit |  | 2 unit
|  | 2 | Trolley with caster |  | 2 unit
1.2. | Accessories | 1 | Pedal foot switch | Pedal foot switch | 2 sets
|  | 2 | Monopolar hand switch with cable (reusable) | Monopolar hand switch with cable (reusable) | 4 sets
|  | 3 | Bipolar cable (reusable) | Bipolar cable (reusable) | 4 sets
|  | 4 | Patient plate, silicon (reusable) | Patient plate, silicon (reusable) | 4 sets
|  | 5 | Patient plate cable (reusable) | Patient plate cable (reusable) | 4 sets
|  | 6 | Electrode set (reusable) | Electrode set (reusable) | 12 sets
|  | 7 | Bipolar forceps (reusable) | Bipolar forceps (reusable) | 12 sets
|  | * | Any other accessories which is required for installation and operation should be included and described on the offered specification.
2. | Technical Specifications | 2.1 | Main Unit | specifications |  | Main Unit | specifications
|  | a) | Mode | Monopolar cutting/coagulation and bipolar
|  | b) | Maximum output
|  |  | - Monopolar cutting | 300W or more | 300W or more
|  |  | - Monopolar coagulation | 200W or more | 200W or more
|  |  | - Bipolar | 100W or more | 100W or more
|  | c) | Monopolar port | 2 or more | 2 or more
|  | d) | Type according to EN 60601-1 | CF | CF
|  | e) | Reusable electrode | Standard set for 10 kinds or more | Standard set for 10 kinds or more
|  | f) | Foot switch | Monopolar and Bipolar | Monopolar and Bipolar
|  | g) | Menu Language | Romanian or Russian
|  | h) | Power requirement | AC 220V±10%, single phase, 50Hz
|  | 1 | Monopolar hand switch with cable (reusable) | Monopolar hand switch with cable (reusable) | 4 sets
|  | 2 | Bipolar cable (reusable) | Bipolar cable (reusable) | 4 sets
|  | 3 | Patient plate, silicon (reusable) | Patient plate, silicon (reusable) | 4 sets
|  | 4 | Patient plate cable (reusable) | Patient plate cable (reusable) | 4 sets
|  | 5 | Electrode set (reusable) | Electrode set (reusable) | 12 sets
|  | 6 | Bipolar forceps (reusable) | Bipolar forceps (reusable) | 12 sets
4. | Consumables | 1 | N/a
* Please add composition, accessories, spare parts, consumables (if applicable)
Note on F6: LL: Quantity is set for 2 units of equipment

## Lot 9
Annex 2
Technical compliance table of the medical equipment for the Medical Service of the Ministry of Internal Affairs of the Republic of Moldova
|  |  | REQUIRED |  |  | OFFERED
Item |  |  |  |  | Quantity | Item |  | Quantity
LOT 9 | OPERATION LIGHT, ULTRALIGHT |  | Composition, accessories, consumables, spare parts |  | 1 | Composition, accessories, consumables, spare parts |  | 1
1.1. | Composition | 1 | Main unit |  | 1 unit
1.2. | Accessories | 1 | Sterilizable handle |  | 1 set
|  | * | Any other accessories which is required for installation and operation should be included and described on the offered specification.
2. | Technical Specifications | 2.1 | Main Unit | specifications |  | Main Unit | specifications
|  | a) | Type | Combination of main and satellite
|  |  |  | ceiling suspension type | ceiling suspension type
|  | b) | Light source | LED | LED
|  | c) | Light intensity at 1 m
|  |  | - Main | 120,000 lux or more (at 1 m) | 120,000 lux or more (at 1 m)
|  |  | - Satellite | 50,000 lux or more (at 1 m) | 50,000 lux or more (at 1 m)
|  | d) | Color Temperature | 4300 K ± 100 K
|  | e) | Life time light source | 40,000 hours or more | 40,000 hours or more
|  | f) | Controls | On the cupola | On the cupola
|  | g) | Movements | Rotation on 3 horizonal axes, Up/Down on 1 axe or more | Rotation on 3 horizonal axes, Up/Down on 1 axe or more
|  | h) | Power requirement | AC 220V±10%, 1 phase, 50Hz
3. | Spare parts | 1 | Sterilizable handle | Pedal foot switch | 2 sets
4. | Consumables | 1 | N/a |  | -
* Please add composition, accessories, spare parts, consumables (if applicable)

## Lot 10
Annex 2
Technical compliance table of the medical equipment for the Medical Service of the Ministry of Internal Affairs of the Republic of Moldova
|  |  | REQUIRED |  |  | OFFERED
Item |  |  |  |  | Quantity | Item |  | Quantity
LOT 10 | ECG, 12-channel |  | Composition, accessories, consumables, spare parts |  | 1 | Composition, accessories, consumables, spare parts |  | 1
1.1. | Composition | 1 | Main unit |  | 1 unit
|  | 2 | Cart | Hand remote control unit | 1 unit
1.2. | Accessories | 1 | Patient cable |  | 1 set
|  | 2 | Patient cable hanger |  | 1 set
|  | 3 | Power cord |  | 1 pc.
|  | 4 | Chest electrode |  | 6 pcs.
|  | 5 | Limb electrode |  | 4 pcs.
|  | 6 | Thermal head cleaner |  | 1 pc.
|  | * | Any other accessories which is required for installation and operation should be included and described on the offered specification.
2. | Technical Specifications | 2.1 | Main unit | specifications |  | Main Unit | specifications
|  | a) | No. of channel | 12 channel
|  | b) | Leads | Standard 12 leads | Standard 12 leads
|  | c) | Sensitivity | 5/10/20 mV or more | 5/10/20 mV or more
|  | d) | Time constant | 3sec or more | 3sec or more
|  | e) | Frequency response | 150 Hz or less | 150 Hz or less
|  | f) | Common mode rejection ratio | 100dB or more | 100dB or more
|  | g) | Maximum electrode potential | ±500mV or more | ±500mV or more
|  | h) | Sampling frequency | 4,000 samples/sec. or more | 4,000 samples/sec. or more
|  | i) | Display |  |  | Display
|  |  | - Device | LCD | LCD
|  |  | - Resolution | 320 x 240 dots or more | 320 x 240 dots or more
|  |  | - Representation | 3/6/12 channels/display | 3/6/12 channels/display
|  | j) | Recorder |  |  | Recorder
|  |  | - Method | Thermal dot array | Thermal dot array
|  |  | - Paper speed | 5, 10, 12.5, 25, 50 mm/sec | 5, 10, 12.5, 25, 50 mm/sec
|  |  | - Paper Size | 210 mm | 210 mm
|  | k) | Power requirement | AC 220V±10%, single phase, 50Hz |  | Power requirement
3. | Spare parts | 1 | Memory card |  | 1 pc.
|  | 2 | Patient cable |  | 1 unit
|  | 3 | Chest electrode |  | 6 pcs.
|  | 4 | Limb electrode |  | 4 pcs.
4. | Consumables | 1 | Recording paper |  | 30 rolls
|  | 2 | Thermal head cleaner |  | 1 pc.
|  | 3 | ECG cream |  | 12 tubes
* Please add composition, accessories, spare parts, consumables (if applicable)

## Lot 11
Annex 2
Technical compliance table of the medical equipment for the Medical Service of the Ministry of Internal Affairs of the Republic of Moldova
|  |  | REQUIRED |  |  | OFFERED
Item |  |  |  |  | Quantity | Item |  | Quantity
LOT 11 | ULTRASOUND Apparatus (General) |  | Composition, accessories, spare parts, consumables |  | 1 | Composition, accessories, spare parts, consumables |  | 1
1.1. | Composition | 1 | Main unit with system control software |  | 1 unit
|  | 2 | Probe A | Probe A | 1 unit
|  | 3 | Probe B | Probe B | 1 unit
|  | 4 | Probe C | Probe C | 1 unit
|  | 5 | Probe D | Probe D | 1 unit
|  | 6 | B/W printer | B/W printer | 1 unit
1.2. | Accessories | 1 | Probe hanger and holder | Probe hanger and holder | 1 set
|  | 2 | Software for measurement | Software for measurement
|  |  | - Cardiac measurement | Cardiac measurement | 1 set
|  |  | (LV, LA, AV, MV, PV or more) | (LV, LA, AV, MV, PV or more)
|  |  | - Obstetric and gynecology measurement | Obstetric and gynecology measurement | 1 set
|  |  | - Vascular measurement (optional) | Vascular measurement (optional) | 1 set
|  | 3 | Software for DICOM (storage and transfer) |  | 1 set
|  | * | Any other accessories which is required for installation and operation should be included and described on the offered specification.
2. | Technical Specifications | 2.1 | Main unit | specifications |  | Main Unit | specifications
|  | a) | Scan mode | Linear, Convex
|  | b) | Monitor | LCD color 17 inch or more | LCD color 17 inch or more
|  | c) | Viewing mode | B-mode, M-mode, Spectrum Doppler, | B-mode, M-mode, Spectrum Doppler,
|  |  |  | Color Doppler(optional), Realtime 3D (optional) or more | Color Doppler(optional), Realtime 3D (optional) or more
|  | d) | Pulse wave Doppler frequency range | 1 kHz or less - 10kHz or more | 1 kHz or less - 10kHz or more
|  | e) | Continuous wave Doppler | 4 kHz or less - 40kHz or more | 4 kHz or less - 40kHz or more
|  | f) | Color Doppler mode | Optional | Optional
|  | g) | Reference signal | Equipped by cardiac unit | Equipped by cardiac unit
|  |  |  | TDI mode or more | TDI mode or more
|  | h) | Number of active probe connector | 3 or more | 3 or more
|  | i) | Power supply | AC 220V±10%, single phase, 50Hz | AC 220V±10%, single phase, 50Hz
|  | 2.2 | Probe A | specifications |  | Probe A | specifications
|  | a) | Type | Convex for abdomen, adult | Convex for abdomen, adult
|  | b) | B-mode | 1.0 - 5.0 MHz or wider | 1.0 - 5.0 MHz or wider
|  | c) | THI mode | Equipped | Equipped
|  | d) | Radius of curvature | 80 mmR or less | 80 mmR or less
|  | 2.3 | Probe B | specifications |  | Probe B | specifications
|  | a) | Type | Cardiac
|  | b) | B-mode | 1.0 - 3.0 MHz or wider
|  | c) | Modes | Flow,PW,CW or more
|  | d) | Scanning angle | 90 ̊ or more
|  | 2.4 | Probe C | specifications |  | Probe C | specifications
|  | a) | Type | Transvaginal
|  | b) | B-mode | 3.0 - 7.0 MHz or wider
|  | c) | THI mode | Equipped
|  | d) | Scanning angle | 180 ̊ or more
|  | 2.5 | Probe D | specifications |  | Probe D | specifications
|  | a) | Type | Linear for surface
|  | b) | B-mode | 3.0 - 11.0MHz or wider
|  | c) | THI mode | Equipped
|  | d) | Scanning width | 38 mm or more
3 | Spare parts | 1 | N/a |  | -
4 | Consumables | 1 | Echo gel (100g or more) |  | 24 tubes
|  | 2 | Printer paper (B/W) |  | 12 rolls
* Please add composition, accessories, spare parts, consumables (if applicable)

## Lot 12
Annex 2
Technical compliance table of the medical equipment for the Medical Service of the Ministry of Internal Affairs of the Republic of Moldova
|  |  | REQUIRED |  |  | OFFERED
Item |  |  |  |  | Quantity | Item |  | Quantity
LOT 12 | LAPAROSCOPE |  | Composition, accessories, spare parts, consumables |  | 1 | Composition, accessories, spare parts, consumables |  | 1
1.1. | Composition | 1 | Laparoscope 30 degrees, 10 mm |  | 1 unit
1.2. | Accessories | 1 | N/a
|  | * | Any other accessories which is required for installation and operation should be included and described on the offered specification.
2. | Technical Specifications | 2.1 | Laparoscope | specifications |  |  | specifications
|  | a) | Compatible with Videocamera GIMMI SNH 890 S.2000.81 II and light cable
3 | Spare parts | 1 | N/a
4 | Consumables | 1 | N/a
* Please add composition, accessories, spare parts, consumables (if applicable)
**The supplier will provide all other additional components if required to complete the installation in the facility indicated by the end-user.

## Lot 13
Annex 2
Technical compliance table of the medical equipment for the Medical Service of the Ministry of Internal Affairs of the Republic of Moldova
|  |  | REQUIRED |  |  | OFFERED
Item |  |  |  |  | Quantity | Item |  | Quantity
LOT 13 | UROLOGY ENDOSCOPE |  | Composition, accessories, spare parts, consumables |  | 1 | Composition, accessories, spare parts, consumables |  | 1
1.1. | Composition | 1 | Urologic Endoscope 12 degrees, 4mm |  | 1 unit
|  | 2 | Urologic Endoscope 30 degrees, 4mm |  | 1 unit
1.2. | Accessories | 1 | N/a
|  | * | Any other accessories which is required for installation and operation should be included and described on the offered specification.
2. | Technical Specifications | 2.1 | Urologic Endoscope | specifications |  |  | specifications
|  | a) | Compatible with KARL STORZ Telecam SL II 20213020 and KARL STORZ Xenon Nova-175 20131520 Light source and light cable
3 | Spare parts | 1 | N/a
4 | Consumables | 1 | N/a
* Please add composition, accessories, spare parts, consumables (if applicable)
**The supplier will provide all other additional components if required to complete the installation in the facility indicated by the end-user.

## Lot 14
Annex 2
Technical compliance table of the medical equipment for the Medical Service of the Ministry of Internal Affairs of the Republic of Moldova
|  |  | REQUIRED |  |  | OFFERED
Item |  |  |  |  | Quantity | Item |  | Quantity
LOT 14 | FIBROGASTRODUODENOSCOPE |  | Composition, accessories, spare parts, consumables |  | 1 | Composition, accessories, spare parts, consumables |  | 1
1.1. | Composition | 1 | Fibrogastroduodenoscope |  | 1 unit
1.2. | Accessories | 1 | N/a
|  | * | Any other accessories which is required for installation and operation should be included and described on the offered specification.
2. | Technical Specifications | 2.1 | Fibrogastroduodenoscope | specifications |  |  | specifications
|  | a) | Compatibile with OLYMPUS CLK-4 light source
3 | Spare parts | 1 | N/a
4 | Consumables | 1 | Loop for cauterization (multiple use) |  | 1 piece
|  | 2 | Biopsy forceps (multiple use) |  | 1 piece
|  | 3 | Injection needle (multiple use) |  | 1 piece
* Please add composition, accessories, spare parts, consumables (if applicable)
**The supplier will provide all other additional components if required to complete the installation in the facility indicated by the end-user.

## Lot 15
Annex 2
Technical compliance table of the medical equipment for the Medical Service of the Ministry of Internal Affairs of the Republic of Moldova
|  |  | REQUIRED |  |  | OFFERED
Item |  |  |  |  | Quantity | Item |  | Quantity
LOT 15 | HIGH PRESSURE STEAM STERILIZER |  | Composition, accessories, spare parts, consumables |  | 1 | Composition, accessories, spare parts, consumables |  | 1
1.1. | Composition | 1 | Main unit |  | 1
|  | 2 | Built-In Digital or Thermal Printer |  | 1 unit
|  | 3 | Built-in Water Treatment System |  | 1 unit
1.2. | Accessories | 1 | Loading Baskets for Sterile Material/Instruments (according to the number of trays) |  | (according to the number of trays)
|  | * | Any other accessories which is required for installation and operation should be included and described on the offered specification.
2. | Technical Specifications | 2.1 | Main Unit | specifications |  | Main Unit | specifications
|  | a) | Door type | 1 door
|  | b) | Chamber volume | 80 to 100 l
|  | c) | Chamber material | Stainless steel
|  | d) | Sterilizing temperature | 121 °C and 134 °C
|  | e) | Safety devices | Equipped
|  | f) | Unloaded Weight (kg) | up to 200 kg
|  | g) | Adjustable feet or wheels
|  | h) | Power requirment | AC 230V or 400V±10%, 3 phase, 50Hz
3 | Spare parts | 1 | Door gasket |  | 1 unit
4 | Consumables | 1 | Printer paper for built-in printer |  | 6 units
|  | 2 | Filter element of filter set |  | 1 unit
* Please add composition, accessories, spare parts, consumables (if applicable)
**The supplier will provide all other additional components if required to complete the installation in the facility indicated by the end-user.

## Lot 16
Annex 2
Technical compliance table of the medical equipment for the Medical Service of the Ministry of Internal Affairs of the Republic of Moldova
|  |  | REQUIRED |  |  | OFFERED
Item |  |  |  |  | Quantity | Item |  | Quantity
LOT 16 | X-RAY EQUIPMENT |  | Composition, accessories, cosumables, spare parts |  | 1 | Composition, accessories, cosumables, spare parts |  | 1
1.1. | Composition | 1 | X-ray generator |  | 1
|  | 2 | Floor-Mounted Tube Support |  | 1
|  | 3 | X Ray tube |  | 1
|  | 4 | Bucky Table |  | 1
|  | 5 | Bucky Stand |  | 1
|  | 6 | Dry Imager |  | 1
|  | 7 | Digital radiography (FPD 14x17 inch) |  | 1
|  | 8 | PACS |  | 1
1.2. | Accessories | 1 | X-ray protective aprons |  | 1
|  | 2 | Interphone with HD video camera and monitor |  | 1
2. | Technical Specifications | 2.1. | X-ray generator | specifications |  | X-ray generator | specifications
|  | a) | Type | Inverter
|  | b) | Capacity | 35 kW or less
|  | c) | X-ray tube connectable number | 1
|  | d) | Max kV | 150 kV or more
|  | e) | Min. kV | 40 kV or less
|  | f) | Max mA | 630 or more
|  | g) | Min mA | 10 mA or less
|  | h) | AEC | Yes
|  | i) | Exposure time | 0,001 sec to 10 sec
|  | j) | Number of phototimer field | 4 field or more
|  | k) | DAP Meter | Yes
|  | l) | Anatomical programs | 400 or more
|  | 2.2. | Floor-Mounted Tube Support | specifications |  | Floor-Mounted Tube Support | specifications
|  | a) | Vertical travel | 1300 mm or more
|  | b) | Longitudinal travel | 1800 mm or more
|  | c) | Transverse travel | 200 mm or more
|  | d) | X Ray Tube rotation on horizontal X axis | CW: 140 degree, CCW: 180 degree
|  | e) | X Ray Tube column rotation on vertyical axis | 180 degree (stretcher position)
|  | f) | Illumination angle meter | Yes
|  | 2.3. | X-Ray tube | specifications |  | X-Ray tube | specifications
|  | a) | Focal spot | 1/2 mm or more | 1/2 mm or more
|  | b) | Max anode heat capacity | 140 kHU or more | 140 kHU or more
|  | c) | Max X-Ray tube voltage | 125 kV or more | 125 kV or more
|  | d) | Collimator | Manual | Manual
|  | 2.4. | Bucky Table | specifications |  | Bucky Table | specifications
|  | a) | Floating table top | Yes (4 way floating table) | Yes (4 way floating table)
|  | b) | Table top dimensions | 800x2300 mm or more | 800x2300 mm or more
|  | c) | Table top height | 700 mm or less | 700 mm or less
|  | d) | Disctance between tabletop and film | 75 mm or less | 75 mm or less
|  | e) | Longitudinal movment | 1100 mm or more | 1100 mm or more
|  | f) | Lateral movment | 250 mm or more | 250 mm or more
|  | g) | Maximum load capacity | 220 kg or more | 220 kg or more
|  | h) | Removable grid | Yes | Yes
|  | 2.5. | Bulky stand | specifications |  | Bulky stand | specifications
|  | a) | Vertical movment | 1300 mm or more
|  | b) | Removable grid | Yes
|  | 2.6. | Dry Imager | Direct thermal |  | Dry Imager | specifications
|  | a) | Printing speed | 80 sheets/hr. or more
|  | b) | Spatial resolution | 320 DPI
|  | c) | Exposure resolution | 14 bit or more
|  | d) | Interface | DICOM
|  | e) | Media outputs | One receive tray or more
|  | f) | Media inputs | Two supply cassettes or more
|  | g) | Media Sizes | 8x10",11x14",14x17" or more
|  | 2.7. | Digital radiography | specifications |  | Digital radiography | specifications
|  | a) | Flat panel detector size | 14x17 or more | 14x17 or more
|  | b) | Scintillator | CSL | CSL
|  | c) | Pixel size | 150 µm or less | 150 µm or less
|  | d) | Bit depth | 14 bits or more | 14 bits or more
|  | e) | Image prevue | 2 sec or less | 2 sec or less
|  | f) | Cycle time (wirless connection) | 20 sec or less | 20 sec or less
|  | g) | Weight | 3,8 kg or less | 3,8 kg or less
|  | h) | Number of exposable images | 200 full images or more/3 h full performance or more | 200 full images or more/3 h full performance or more
|  | i) | DICOM | MPPS,MWM,Storage,Print | MPPS,MWM,Storage,Print
|  | j) | Data interface/Detector | WIGI/Ethernet | WIGI/Ethernet
|  | 2.8. | PACS | specifications |  | PACS | specifications
|  | a) | HDD | 1 TB or more, RAID 10
|  | b) | Processor | Intel Quad-Core Xeon X3450 CPU(2.66GHz) or similar
|  | c) | RAM | 4 GB or more
|  | d) | Software | Cine Tools, Measurements, Annotations, DICOM send, CD, administration and remote diagnostics
|  | e) | Monitor | Medical Grade, 2 MP or more, 20" or more
|  | f) | WEB or LAN Access Licenses | 3 or more
|  | g) | Interconnection and Set up works* | All works for interconnection of the existing Xray devices in the institution with the PACS System will be supported by the supplier
|  | 2.9. | X-ray protective aprons | specifications |  | X-ray protective aprons | specifications
|  | a) | Lead equivalent | 0.35 mmPb
|  | b) | Size | L
|  | 2.10 | Interphone with HD video camera and monitor | specifications |  | Interphone with HD video camera and monitor | specifications
|  | a) | HD Monitor | 14 inch or more
|  | b) | Camera | HD resolution or higher
|  | c) | Interphone | Bidirectional, wall mounted, button controlled
|  | 2.11 | Power requirement | specifications |  | Power requirement | specifications
|  | a) | AC 240V±10%,two phase, 50Hz and AC 220V±10%, single phase, 50Hz
3. | Spare Parts | 1 | N/a
4 | Consumables | 1 | Recording film (8 x 10 inch) |  | 200 sheets
|  | 2 | Recording film (11 x 14 inch) |  | 200 sheets
|  | 3 | Recording film (14 x 17 inch) |  | 200 sheets