## Sheet1
SPECIFICATION SHEET
Brand/Model
Code No. | Vacuum-assisted fetal delivery device |  |  | Quantity: | 8
Purpose : | The device is used to assist delivery of a baby.
please, fill in below YES or NO towards the required specs.
1. | Composition
1-1 | Composition
| Vacuum-assisted fetal delivery device | : | 8 units
1-2 | Accessories
| Silicon Cups set for fetus vacuum extraction | : | 40 – 60 mm
Any other accessory which is required for installation and operation should be included and described on your offered specification
2. | Specifications
─ | Maximum flow | : | 30 L/min or more
─ | Maximum vacuum | : | 60 kPa or more
─ | Noise level | : | 60 dB or less
─ | Voltage | : | ±5%, calculated at 230V /60 Hz
─ | Collection jars (autoclavable) | : | ≥1 L
─ | Selector knob for double collection jar exchange | : | yes
─ | Trolley | : | yes
─ | Hydrophobic PTFE bacteria filter | : | 99,95% particle retention and 99,99% bacterial challenge efficiency
─ | On/Off Switch | : | yes
─ | High Vacuum Gauge including a display | : | 0-100kPa / 0-760 mmHg
─ | On/Off foot-operated pedal | : | yes
─ | Suction pressure-regulating knob | : | yes
─ | PVC suction tube length | : | 3m or more
─ | Power | : | 110 W or more
3. | General Requirements
Compliant with MDD 93/42/EEC or US FDA Regulations
Manufacturer ISO 13485
Warranty at least 36 months
4. | Training
Training for the personnel who will be using and maintaining the equipment
Training in Russian language
5. | Documentation
User’s manuals, preferable in Russian language (optional English)
Maintenance manual, preferable in Russian language (optional English)